## Sayfa1
Laboratory errors | Analyte | Effects on test results | Reference | Reference 2
Hemolysis | Chloride (serum) | Falsely decreased Chloride | Lippi G, Salvagno GL, Montagnana M, Brocco G, Guidi GC. Influence of hemolysis on routine clinical chemistry testing. Clin Chem Lab Med. 2006;44(3):311-6.
Hemolysis | GGT (Gamma-glutamyltransferase) (serum) | Falsely decreased GGT | Knezevic CE, Ness MA, Tsang PHT, Tenney BJ, Marzinke MA. Establishing hemolysis and lipemia acceptance thresholds for clinical chemistry tests. Clin Chim Acta. 2020;510:459-65.
Hemolysis | Bicarbonate (HCO3-) (blood) | Falsely decreased Bicarbonate | Knezevic CE, Ness MA, Tsang PHT, Tenney BJ, Marzinke MA. Establishing hemolysis and lipemia acceptance thresholds for clinical chemistry tests. Clin Chim Acta. 2020;510:459-65.
Hemolysis | Direct bilirubin (serum) | Falsely decreased Direct Bilirubin | Knezevic CE, Ness MA, Tsang PHT, Tenney BJ, Marzinke MA. Establishing hemolysis and lipemia acceptance thresholds for clinical chemistry tests. Clin Chim Acta. 2020;510:459-65.
Hemolysis | ALP (alkaline phosphatase) (serum) | Falsely decreased Alkaline Phosphatase (ALP) | Agarwal S, Vargas G, Nordstrom C, Tam E, Buffone GJ, Devaraj S. Effect of interference from hemolysis, icterus and lipemia on routine pediatric clinical chemistry assays. Clin Chim Acta. 2015;438:241-5.
Hemolysis | Potassium (serum) | Falsely increased Potassium | Knezevic CE, Ness MA, Tsang PHT, Tenney BJ, Marzinke MA. Establishing hemolysis and lipemia acceptance thresholds for clinical chemistry tests. Clin Chim Acta. 2020;510:459-65.
Hemolysis | AST (Aspartate aminotransferase) (serum) | Falsely increased AST | Knezevic CE, Ness MA, Tsang PHT, Tenney BJ, Marzinke MA. Establishing hemolysis and lipemia acceptance thresholds for clinical chemistry tests. Clin Chim Acta. 2020;510:459-65.
Hemolysis | CK (Creatine Kinase) (serum) | Falsely increased CK | Knezevic CE, Ness MA, Tsang PHT, Tenney BJ, Marzinke MA. Establishing hemolysis and lipemia acceptance thresholds for clinical chemistry tests. Clin Chim Acta. 2020;510:459-65.
Hemolysis | LDH (Lactate dehydrogenase) (serum) | Falsely increased LDH | Knezevic CE, Ness MA, Tsang PHT, Tenney BJ, Marzinke MA. Establishing hemolysis and lipemia acceptance thresholds for clinical chemistry tests. Clin Chim Acta. 2020;510:459-65.
Hemolysis | Iron (serum) | Falsely increased Iron | Knezevic CE, Ness MA, Tsang PHT, Tenney BJ, Marzinke MA. Establishing hemolysis and lipemia acceptance thresholds for clinical chemistry tests. Clin Chim Acta. 2020;510:459-65.
Hemolysis | Lipase (serum) | Falsely increased Lipase | Knezevic CE, Ness MA, Tsang PHT, Tenney BJ, Marzinke MA. Establishing hemolysis and lipemia acceptance thresholds for clinical chemistry tests. Clin Chim Acta. 2020;510:459-65.
Hemolysis | Total protein (serum) | Falsely increased Total Protein | Knezevic CE, Ness MA, Tsang PHT, Tenney BJ, Marzinke MA. Establishing hemolysis and lipemia acceptance thresholds for clinical chemistry tests. Clin Chim Acta. 2020;510:459-65.
Hemolysis | Phosphorus (inorganic) (serum) | Falsely increased Phosphorus | Glick MR, Ryder KW, Jackson SA. Graphical comparisons of interferences in clinical chemistry instrumentation. Clin Chem. 1986;32(3):470-5.
Hemolysis | Magnesium (serum) | Falsely increased Magnesium | Lippi G, Salvagno GL, Montagnana M, Brocco G, Guidi GC. Influence of hemolysis on routine clinical chemistry testing. Clin Chem Lab Med. 2006;44(3):311-6.
Hemolysis | Albumin (serum) | Falsely increased Albumin | Agarwal S, Vargas G, Nordstrom C, Tam E, Buffone GJ, Devaraj S. Effect of interference from hemolysis, icterus and lipemia on routine pediatric clinical chemistry assays. Clin Chim Acta. 2015;438:241-5.
Hemolysis | Ammonia (plasma) | Falsely increased Ammonia | Agarwal S, Vargas G, Nordstrom C, Tam E, Buffone GJ, Devaraj S. Effect of interference from hemolysis, icterus and lipemia on routine pediatric clinical chemistry assays. Clin Chim Acta. 2015;438:241-5.
Hemolysis | Triglyceride (serum) | Falsely increased Triglyceride | Knezevic CE, Ness MA, Tsang PHT, Tenney BJ, Marzinke MA. Establishing hemolysis and lipemia acceptance thresholds for clinical chemistry tests. Clin Chim Acta. 2020;510:459-65.
Hemolysis | Total bilirubin (serum) | Falsely increased or decreased total bilirubin | Agarwal S, Vargas G, Nordstrom C, Tam E, Buffone GJ, Devaraj S. Effect of interference from hemolysis, icterus and lipemia on routine pediatric clinical chemistry assays. Clin Chim Acta. 2015;438:241-5. | Uçar KT, Çat A, Gümüş A, Nurlu N. Interferograms plotted with reference change value (RCV) may facilitate the management of hemolyzed samples. J Med Biochem. 2022;41(1):53-61.
Hemolysis | ALT (Alanine aminotransferase) (serum) | Falsely increased or decreased ALT | Agarwal S, Vargas G, Nordstrom C, Tam E, Buffone GJ, Devaraj S. Effect of interference from hemolysis, icterus and lipemia on routine pediatric clinical chemistry assays. Clin Chim Acta. 2015;438:241-5. | Uçar KT, Çat A, Gümüş A, Nurlu N. Interferograms plotted with reference change value (RCV) may facilitate the management of hemolyzed samples. J Med Biochem. 2022;41(1):53-61.
Icterus | Fructosamine (serum) | Falsely increased Fructosamine | Nicolay A, Lorec AM, Gomez G, Portugal H. Icteric human samples: Icterus index and method of estimating an interference-free value for 16 biochemical analyses. J Clin Lab Anal. 2018;32(2).
Icterus | Phosphorus (inorganic) (serum) | Falsely increased Phosphorous | Nicolay A, Lorec AM, Gomez G, Portugal H. Icteric human samples: Icterus index and method of estimating an interference-free value for 16 biochemical analyses. J Clin Lab Anal. 2018;32(2).
Icterus | Creatinine (serum) | Falsely decreased Creatinine | Nicolay A, Lorec AM, Gomez G, Portugal H. Icteric human samples: Icterus index and method of estimating an interference-free value for 16 biochemical analyses. J Clin Lab Anal. 2018;32(2).
Icterus | HDL Cholesterol (serum) | Falsely decreased HDL Cholesterol | Nicolay A, Lorec AM, Gomez G, Portugal H. Icteric human samples: Icterus index and method of estimating an interference-free value for 16 biochemical analyses. J Clin Lab Anal. 2018;32(2).
Icterus | Total cholesterol (serum) | Falsely decreased Total Cholesterol | Nicolay A, Lorec AM, Gomez G, Portugal H. Icteric human samples: Icterus index and method of estimating an interference-free value for 16 biochemical analyses. J Clin Lab Anal. 2018;32(2).
Icterus | Total protein (serum) | Falsely decreased Total Protein | Nicolay A, Lorec AM, Gomez G, Portugal H. Icteric human samples: Icterus index and method of estimating an interference-free value for 16 biochemical analyses. J Clin Lab Anal. 2018;32(2).
Icterus | Triglyceride (serum) | Falsely decreased Tryglyceride | Nicolay A, Lorec AM, Gomez G, Portugal H. Icteric human samples: Icterus index and method of estimating an interference-free value for 16 biochemical analyses. J Clin Lab Anal. 2018;32(2).
Icterus | Uric acid (serum) | Falsely decreased Uric Acid | Nicolay A, Lorec AM, Gomez G, Portugal H. Icteric human samples: Icterus index and method of estimating an interference-free value for 16 biochemical analyses. J Clin Lab Anal. 2018;32(2).
Icterus | ADA (Adenosine deaminase activity) Body Fluid | Falsely decreased ADA | Allison TR, Hunsaker JJH, La'ulu SL, Genzen JR. Evaluation of an adenosine deaminase (ADA) assay in serum, pleural, pericardial, peritoneal, and cerebrospinal fluids on the Roche cobas c501 analyzer. Clin Biochem. 2022;109-110:57-63.
Lipemia | HbA1c (Hemoglobin A1c) (blood) | Falsely increased HbA1c (Hemoglobin A1c) | Yıldırım Demirçin E, Tutar S, Sezer S, Akbıyık F, Yılmaz G. Lipemia impact on HBA1C measurements in patients with and without hemoglobin variant: a rational laboratory use evaluation. Scand J Clin Lab Invest. 2023;83(3):194-9.
Lipemia | Sodium (serum) | Falsely decreased Sodium | Barkhuizen M, Hoffmann M, Zöllner EW, Erasmus RT, Zemlin AE. Case report: An index of suspicion in hyponatraemia. Biochem Med (Zagreb). 2019;29(1):011002.
Macroenzymes | Amylase, total (serum) | Falsely increased Amylase | Čásenská J, Franeková J, Mačinga P, Jabor A. Significant elevations of serum amylase caused by macroamylase: Case reports and detection possibilities. J Clin Lab Anal. 2023;37(5):e24859.
Macroenzymes | AST (Aspartate aminotransferase) (serum) | Falsely increased AST | Can Ü, Sayar E, Serdar MA. Macro-aspartate aminotransferase (AST): a case report / Makro aspartat aminotransferaz (AST): Bir olgu takdimi. Turkish Journal of Biochemistry. 2016;41(1):55-7.
Macroenzymes | CK (Creatine Kinase) (serum) | Falsely increased CK | Wyness SP, Hunsaker JJ, La'ulu SL, Rao LV, Roberts WL. Detection of macro-creatine kinase and macroamylase by polyethylene glycol precipitation and ultrafiltration methods. Clin Chim Acta. 2011;412(23-24):2052-7.
Heterophilic antibodies | hs-Troponin I&T (high sensitive) (serum) | Falsely elevated hs-Troponin I&T (high sensitive) | Chaulin AM. False-Positive Causes in Serum Cardiac Troponin Levels. J Clin Med Res. 2022;14(2):80-7.
Heterophilic antibodies | CK-MB (Creatine kinase-MB) (mass) (serum) | Falsely elevated CK-MB (mass) | Sosolik RC, Hitchcock CL, Becker WJ. Heterophilic antibodies produce spuriously elevated concentrations of the MB isoenzyme of creatine kinase in a selected patient population. Am J Clin Pathol. 1997;107(5):506-10.
Heterophilic antibodies | Myoglobin (serum) | Falsely elevated Myoglobin | Tate J, Ward G. Interferences in immunoassay. Clin Biochem Rev. 2004;25(2):105-20.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | LDL Cholesterol (serum) | Falsely increased LDL Cholesterol | Zheng C, Zhou W, Zhou R, Zhang H, Liang S, Zhang W, et al. The necessity for improving lipid testing reagents: A real world study. Clin Chim Acta. 2023;548:117529.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | UIBC (Iron Binding Capacity, unsaturated) (serum) | Falsely increased UIBC (Iron Binding Capacity, unsaturated) | Ikuta K, Ito S, Tanaka H, Sasaki K, Torimoto Y, Fujiya M, et al. Interference of deferasirox with assays for serum iron and serum unsaturated iron binding capacity during iron chelating therapy. Clin Chim Acta. 2011;412(23-24):2261-6.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | Blood Urea Nitrogen (BUN) (serum) | Falsely increased Blood Urea Nitrogen (BUN) | Luna-Záizar H, Virgen-Montelongo M, Cortez-Álvarez CR, Ruiz-Quezada SL, Escutia-Gutiérrez R, García-Lemus CR, et al. In vitro interference by acetaminophen, aspirin, and metamizole in serum measurements of glucose, urea, and creatinine. Clin Biochem. 2015;48(7-8):538-41
Biotin interference | hs-Troponin I&T (high sensitive) (serum) | Falsely decreased hs-Troponin I&T (high sensitive) | Willeman T, Casez O, Faure P, Gauchez AS. Evaluation of biotin interference on immunoassays: new data for troponin I, digoxin, NT-Pro-BNP, and progesterone. Clin Chem Lab Med. 2017;55(10):e226-e9.
Biotin interference | BNP and NT-ProBNP (NT-Pro B-Type Natriuretic Peptide) (serum, plasma) | Falsely decreased BNP and NT-ProBNP | Willeman T, Casez O, Faure P, Gauchez AS. Evaluation of biotin interference on immunoassays: new data for troponin I, digoxin, NT-Pro-BNP, and progesterone. Clin Chem Lab Med. 2017;55(10):e226-e9.
Late centrifugation after sample collection | Magnesium (serum) | Falsely elevated Magnesium | Julia O, Wan Azman WN, Tan Say K, Abdul Wahab NA, Samantha X-Y, Law Xi L, et al. Effects of Time Delay in Processing Common Clinical Biochemical Parameters in an Accredited Laboratory. IIUM Medical Journal Malaysia. 2022;21(4).
White particulate matter consisted of fibrinogen, platelets and leukocytes and the fatty droplets | Glucose (serum) | Falsely decreased Glucose | Cadamuro J, Wiedemann H, Felder TK, Mrazek C, Kipman U, Hannes O, et al. What´s floating on my plasma? Biochem Med (Zagreb). 2017;27(2):430-3.
Anti-troponin antibodies | hs-Troponin I&T (high sensitive) (serum) | Falsely elevated hs-Troponin I&T (high sensitive) | Favresse J, Bayart JL, Gruson D, Bernardini S, Clerico A, Perrone M. The underestimated issue of non-reproducible cardiac troponin I and T results: case series and systematic review of the literature. Clin Chem Lab Med. 2021;59(7):1201-11.
IgM antibody response to the E. coli infection | Troponin I (serum) | Falsely elevated Troponin I | Covinsky M, Laterza O, Pfeifer JD, Farkas-Szallasi T, Scott MG. An IgM lambda antibody to Escherichia coli produces false-positive results in multiple immunometric assays. Clin Chem. 2000;46(8 Pt 1):1157-61.
Contamination with detergents, fat emulsions, or samples with methotrexate | Phosphorus (inorganic) (serum) | Falsely elevated inorganic Phosphorus | Wu AHB. Tietz Clinical Guide to Laboratory Tests. 4 ed: Saunders/Elsevier; 2006.
Contamination of distilled water with calcium salts | Calcium (serum) | Falsely elevated Calcium | Wu AHB. Tietz Clinical Guide to Laboratory Tests. 4 ed: Saunders/Elsevier; 2006.
Intake of Iodixanol contrast agent | Calcium (serum) | Falsely decreased Calcium | Otnes S, Fogh-Andersen N, Rømsing J, Thomsen HS. Analytical Interference by Contrast Agents in Biochemical Assays. Contrast Media Mol Imaging. 2017;2017:1323802.
Intake of drugs prepared with bromide | Chloride (serum) | Falsely elevated Chloride | Elin RJ, Robertson EA, Johnson E. Bromide interferes with determination of chloride by each of four methods. Clin Chem. 1981;27(5):778-9.
Sample contamination with heavy metals, borate, usage of some serum separators | Pseudocholinesterase (serum) | Falsely decreased Pseudocholinesterase | Wu AHB. Tietz Clinical Guide to Laboratory Tests. 4 ed: Saunders/Elsevier; 2006.
Hemolysis | Folic acid (serum) | Falsely elevated Folic acid | Wu AHB. Tietz Clinical Guide to Laboratory Tests. 4 ed: Saunders/Elsevier; 2006.
Hemolysis | Insulin (serum) | Falsely decreased Insulin | Wu AHB. Tietz Clinical Guide to Laboratory Tests. 4 ed: Saunders/Elsevier; 2006.
Hemolysis | C-peptide (serum) | Falsely decreased C-peptide | Wu AHB. Tietz Clinical Guide to Laboratory Tests. 4 ed: Saunders/Elsevier; 2006.
Hemolysis | Erythrocyte Sedimentation Rate (blood) | Falsely decreased Erythrocyte Sedimentation Rate | Wu AHB. Tietz Clinical Guide to Laboratory Tests. 4 ed: Saunders/Elsevier; 2006.
Icterus | Procalcitonin (serum) | Falsely increased Procalcitonin | Pagaduan JV, Tam E, Devaraj S. Validation of the Procalcitonin Assay on the Abbott Architect i1000. J Appl Lab Med. 2019;3(6):936-42.
Lipemia | Ferritin (serum) | Falsely elevated Ferritin | Agarwal S, Vargas G, Nordstrom C, Tam E, Buffone GJ, Devaraj S. Effect of interference from hemolysis, icterus and lipemia on routine pediatric clinical chemistry assays. Clin Chim Acta. 2015;438:241-5.
Lipemia | 25-OH Vit D (25-hydroxyvitamin D) (serum) | Falsely decreased Vitamin D | Agarwal S, Vargas G, Nordstrom C, Tam E, Buffone GJ, Devaraj S. Effect of interference from hemolysis, icterus and lipemia on routine pediatric clinical chemistry assays. Clin Chim Acta. 2015;438:241-5.
Biotin interference | 25-OH Vit D (25-hydroxyvitamin D) (serum) | Falsely elevated 25 OH Vitamin D | Piketty ML, Prie D, Sedel F, Bernard D, Hercend C, Chanson P, et al. High-dose biotin therapy leading to false biochemical endocrine profiles: validation of a simple method to overcome biotin interference. Clin Chem Lab Med. 2017;55(6):817-25.
Biotin interference | Vitamin B12 (serum) | Falsely elevated Vitamin B12 | Trambas C, Lu Z, Yen T, Sikaris K. Characterization of the scope and magnitude of biotin interference in susceptible Roche Elecsys competitive and sandwich immunoassays. Ann Clin Biochem. 2018;55(2):205-15.
Biotin interference | Cortisol (serum) | Falsely elevated Cortisol | Piketty ML, Prie D, Sedel F, Bernard D, Hercend C, Chanson P, et al. High-dose biotin therapy leading to false biochemical endocrine profiles: validation of a simple method to overcome biotin interference. Clin Chem Lab Med. 2017;55(6):817-25.
Biotin interference | Anti-TG (thyroglobulin antibody) (serum) | Falsely elevated Anti-Tg | Trambas C, Lu Z, Yen T, Sikaris K. Characterization of the scope and magnitude of biotin interference in susceptible Roche Elecsys competitive and sandwich immunoassays. Ann Clin Biochem. 2018;55(2):205-15.
Biotin interference | Anti-TPO (Thyroid peroxidase antibody) (serum) | Falsely elevated Anti-TPO | Trambas C, Lu Z, Yen T, Sikaris K. Characterization of the scope and magnitude of biotin interference in susceptible Roche Elecsys competitive and sandwich immunoassays. Ann Clin Biochem. 2018;55(2):205-15.
Biotin interference | Anti-TSHR (anti-TSH receptor) | Falsely elevated Anti-TSHR | Trambas C, Lu Z, Yen T, Sikaris K. Characterization of the scope and magnitude of biotin interference in susceptible Roche Elecsys competitive and sandwich immunoassays. Ann Clin Biochem. 2018;55(2):205-15.
Biotin interference | FT3 (free triiodothyronine) (serum) | Falsely elevated Free T3 | Piketty ML, Prie D, Sedel F, Bernard D, Hercend C, Chanson P, et al. High-dose biotin therapy leading to false biochemical endocrine profiles: validation of a simple method to overcome biotin interference. Clin Chem Lab Med. 2017;55(6):817-25.
Biotin interference | FT4 (free thyroxine) (serum) | Falsely elevated Free T4 | Piketty ML, Prie D, Sedel F, Bernard D, Hercend C, Chanson P, et al. High-dose biotin therapy leading to false biochemical endocrine profiles: validation of a simple method to overcome biotin interference. Clin Chem Lab Med. 2017;55(6):817-25.
Biotin interference | DHEA-S (dehydroepiandrosterone sulfate) (serum) | Falsely elevated DHEA-S | Piketty ML, Prie D, Sedel F, Bernard D, Hercend C, Chanson P, et al. High-dose biotin therapy leading to false biochemical endocrine profiles: validation of a simple method to overcome biotin interference. Clin Chem Lab Med. 2017;55(6):817-25.
Biotin interference | Progesterone (serum) | Falsely elevated Progesteron | Piketty ML, Prie D, Sedel F, Bernard D, Hercend C, Chanson P, et al. High-dose biotin therapy leading to false biochemical endocrine profiles: validation of a simple method to overcome biotin interference. Clin Chem Lab Med. 2017;55(6):817-25.
Biotin interference | Testosterone (serum) | Falsely elevated Testosterone | Piketty ML, Prie D, Sedel F, Bernard D, Hercend C, Chanson P, et al. High-dose biotin therapy leading to false biochemical endocrine profiles: validation of a simple method to overcome biotin interference. Clin Chem Lab Med. 2017;55(6):817-25.
Biotin interference | Digoxin (serum) | Falsely elevated Digoxin | Piketty ML, Prie D, Sedel F, Bernard D, Hercend C, Chanson P, et al. High-dose biotin therapy leading to false biochemical endocrine profiles: validation of a simple method to overcome biotin interference. Clin Chem Lab Med. 2017;55(6):817-25.
Biotin interference | Folic acid (serum) | Falsely elevated Folate | Piketty ML, Prie D, Sedel F, Bernard D, Hercend C, Chanson P, et al. High-dose biotin therapy leading to false biochemical endocrine profiles: validation of a simple method to overcome biotin interference. Clin Chem Lab Med. 2017;55(6):817-25.
Biotin interference | Total-Testosterone (serum) | Falsely elevated Total testosterone | Bizzarri C, Antonella Giannone G, Gervasoni J, Benedetti S, Albanese F, Dello Strologo L, et al. Unusual Presentation of Denys-Drash Syndrome in a Girl with Undisclosed Consumption of Biotin. J Clin Res Pediatr Endocrinol. 2021;13(3):347-52.
Biotin interference | Aldosterone (serum) | Falsely elevated Aldosterone | Knudsen CS, Adelborg K, Søndergaard E, Parkner T. Biotin interference in routine IDS-iSYS immunoassays for aldosterone, renin, insulin-like growth factor 1, growth hormone and bone alkaline phosphatase. Scand J Clin Lab Invest. 2022;82(1):6-11.
Biotin interference | Estradiol (serum) | Falsely elevated Estradiol | Piketty ML, Prie D, Sedel F, Bernard D, Hercend C, Chanson P, et al. High-dose biotin therapy leading to false biochemical endocrine profiles: validation of a simple method to overcome biotin interference. Clin Chem Lab Med. 2017;55(6):817-25.
Biotin interference | PTH (Parathyroid hormone) (serum) | Falsely decreased PTH | Trambas C, Lu Z, Yen T, Sikaris K. Characterization of the scope and magnitude of biotin interference in susceptible Roche Elecsys competitive and sandwich immunoassays. Ann Clin Biochem. 2018;55(2):205-15.
Biotin interference | Renin (plasma) | Falsely decreased Renin | Knudsen CS, Adelborg K, Søndergaard E, Parkner T. Biotin interference in routine IDS-iSYS immunoassays for aldosterone, renin, insulin-like growth factor 1, growth hormone and bone alkaline phosphatase. Scand J Clin Lab Invest. 2022;82(1):6-11.
Biotin interference | SHBG (Sex hormone binding globulin) (serum) | Falsely decreased SHBG | Trambas C, Lu Z, Yen T, Sikaris K. Characterization of the scope and magnitude of biotin interference in susceptible Roche Elecsys competitive and sandwich immunoassays. Ann Clin Biochem. 2018;55(2):205-15.
Biotin interference | B-HCG (Beta-Human chorionic gonadotropin) (serum) | Falsely decreased Beta-HCG | Trambas C, Lu Z, Yen T, Sikaris K. Characterization of the scope and magnitude of biotin interference in susceptible Roche Elecsys competitive and sandwich immunoassays. Ann Clin Biochem. 2018;55(2):205-15.
Biotin interference | TSH (Thyroid stimulating hormone) (serum) | Falsely decreased TSH | Stieglitz HM, Korpi-Steiner N, Katzman B, Mersereau JE, Styner M. Suspected Testosterone-Producing Tumor in a Patient Taking Biotin Supplements. J Endocr Soc. 2018;2(6):563-9.
Biotin interference | Insulin (serum) | Falsely decreased Insulin | Trambas C, Lu Z, Yen T, Sikaris K. Characterization of the scope and magnitude of biotin interference in susceptible Roche Elecsys competitive and sandwich immunoassays. Ann Clin Biochem. 2018;55(2):205-15.
Biotin interference | C-peptide (serum) | Falsely decreased C-peptide | Trambas C, Lu Z, Yen T, Sikaris K. Characterization of the scope and magnitude of biotin interference in susceptible Roche Elecsys competitive and sandwich immunoassays. Ann Clin Biochem. 2018;55(2):205-15.
Biotin interference | Troponin T (serum) | Falsely decreased Troponin T | Trambas C, Lu Z, Yen T, Sikaris K. Characterization of the scope and magnitude of biotin interference in susceptible Roche Elecsys competitive and sandwich immunoassays. Ann Clin Biochem. 2018;55(2):205-15.
Biotin interference | ACTH (Adrenocorticotropic Hormone) (plasma) | Falsely decreased ACTH | Trambas C, Lu Z, Yen T, Sikaris K. Characterization of the scope and magnitude of biotin interference in susceptible Roche Elecsys competitive and sandwich immunoassays. Ann Clin Biochem. 2018;55(2):205-15.
Biotin interference | Free PSA (Prostate specific antigen) (serum) | Falsely decreased free PSA | Trambas C, Lu Z, Yen T, Sikaris K. Characterization of the scope and magnitude of biotin interference in susceptible Roche Elecsys competitive and sandwich immunoassays. Ann Clin Biochem. 2018;55(2):205-15.
Biotin interference | Total PSA (Prostate specific antigen) (serum) | Falsely decreased total PSA | Trambas C, Lu Z, Yen T, Sikaris K. Characterization of the scope and magnitude of biotin interference in susceptible Roche Elecsys competitive and sandwich immunoassays. Ann Clin Biochem. 2018;55(2):205-15.
Biotin interference | CA 19-9 (cancer antigen 19-9) (serum) | Falsely decreased CA 19-9 | Trambas C, Lu Z, Yen T, Sikaris K. Characterization of the scope and magnitude of biotin interference in susceptible Roche Elecsys competitive and sandwich immunoassays. Ann Clin Biochem. 2018;55(2):205-15.
Biotin interference | CEA (Carcinoembryonic antigen) (serum) | Falsely decreased CEA | Trambas C, Lu Z, Yen T, Sikaris K. Characterization of the scope and magnitude of biotin interference in susceptible Roche Elecsys competitive and sandwich immunoassays. Ann Clin Biochem. 2018;55(2):205-15.
Biotin interference | CA 125 (cancer antigen 125) (serum) | Falsely decreased CA 125 | Trambas C, Lu Z, Yen T, Sikaris K. Characterization of the scope and magnitude of biotin interference in susceptible Roche Elecsys competitive and sandwich immunoassays. Ann Clin Biochem. 2018;55(2):205-15.
Biotin interference | CA 15-3 (cancer antigen 15-3) (serum) | Falsely decreased CA15-3 | Trambas C, Lu Z, Yen T, Sikaris K. Characterization of the scope and magnitude of biotin interference in susceptible Roche Elecsys competitive and sandwich immunoassays. Ann Clin Biochem. 2018;55(2):205-15.
Biotin interference | AFP (Alpha feto protein) (serum) | Falsely decreased AFP | Trambas C, Lu Z, Yen T, Sikaris K. Characterization of the scope and magnitude of biotin interference in susceptible Roche Elecsys competitive and sandwich immunoassays. Ann Clin Biochem. 2018;55(2):205-15.
Biotin interference | FSH (follicle stimulating hormone) (serum) | Falsely increase or decreased FSH | Piketty ML, Prie D, Sedel F, Bernard D, Hercend C, Chanson P, et al. High-dose biotin therapy leading to false biochemical endocrine profiles: validation of a simple method to overcome biotin interference. Clin Chem Lab Med. 2017;55(6):817-25.
Biotin interference | LH (Luteinizing hormone) (serum) | Falsely increase or decreased LH (28) | Piketty ML, Prie D, Sedel F, Bernard D, Hercend C, Chanson P, et al. High-dose biotin therapy leading to false biochemical endocrine profiles: validation of a simple method to overcome biotin interference. Clin Chem Lab Med. 2017;55(6):817-25.
Biotin interference | Prolactin (serum) | Falsely increase or decreased Prolactin | Piketty ML, Prie D, Sedel F, Bernard D, Hercend C, Chanson P, et al. High-dose biotin therapy leading to false biochemical endocrine profiles: validation of a simple method to overcome biotin interference. Clin Chem Lab Med. 2017;55(6):817-25.
Heterophilic antibodies | TSH (Thyroid stimulating hormone) (serum) | Falsely elevated TSH | Tate J, Ward G. Interferences in immunoassay. Clin Biochem Rev. 2004;25(2):105-20.
Heterophilic antibodies | FT3 (free triiodothyronine) (serum) | Falsely elevated Free T3 | Favresse J, Burlacu MC, Maiter D, Gruson D. Interferences With Thyroid Function Immunoassays: Clinical Implications and Detection Algorithm. Endocr Rev. 2018;39(5):830-50.
Heterophilic antibodies | FT4 (free thyroxine) (serum) | Falsely elevated Free T4 | Favresse J, Burlacu MC, Maiter D, Gruson D. Interferences With Thyroid Function Immunoassays: Clinical Implications and Detection Algorithm. Endocr Rev. 2018;39(5):830-50.
Heterophilic antibodies | Thyroglobulin (serum) | Falsely elevated Thyroglobulin | Tate J, Ward G. Interferences in immunoassay. Clin Biochem Rev. 2004;25(2):105-20.
Heterophilic antibodies | ACTH (Adrenocorticotropic Hormone) (plasma) | Falsely elevated ACTH | Donegan DM, Algeciras-Schimnich A, Hamidi O, Young WF, Nippoldt T, Bancos I, et al. Corticotropin hormone assay interference: A case series. Clin Biochem. 2019;63:143-7.
Heterophilic antibodies | FSH (follicle stimulating hormone) (serum) | Falsely elevated FSH | Gulbahar O, Konca Degertekin C, Akturk M, Yalcin MM, Kalan I, Atikeler GF, et al. A Case With Immunoassay Interferences in the Measurement of Multiple Hormones. J Clin Endocrinol Metab. 2015;100(6):2147-53.
Heterophilic antibodies | PTH (Parathyroid hormone) (serum) | Falsely elevated PTH | Gulbahar O, Konca Degertekin C, Akturk M, Yalcin MM, Kalan I, Atikeler GF, et al. A Case With Immunoassay Interferences in the Measurement of Multiple Hormones. J Clin Endocrinol Metab. 2015;100(6):2147-53.
Heterophilic antibodies | Insulin-like growth factor 1 (IGF-1) | Falsely elevated IGF-1 | Gulbahar O, Konca Degertekin C, Akturk M, Yalcin MM, Kalan I, Atikeler GF, et al. A Case With Immunoassay Interferences in the Measurement of Multiple Hormones. J Clin Endocrinol Metab. 2015;100(6):2147-53.
Heterophilic antibodies | Prolactin (serum) | Falsely elevated prolactin | Gulbahar O, Konca Degertekin C, Akturk M, Yalcin MM, Kalan I, Atikeler GF, et al. A Case With Immunoassay Interferences in the Measurement of Multiple Hormones. J Clin Endocrinol Metab. 2015;100(6):2147-53.
Heterophilic antibodies | B-HCG (Beta-Human chorionic gonadotropin) (serum) | Falsely elevated B-HCG | Gulbahar O, Konca Degertekin C, Akturk M, Yalcin MM, Kalan I, Atikeler GF, et al. A Case With Immunoassay Interferences in the Measurement of Multiple Hormones. J Clin Endocrinol Metab. 2015;100(6):2147-53.
Heterophilic antibodies | Erythropoietin (serum) | Falsely elevated Erythropoietin | Babushok DV, Cardamone D, Rulander N, Master SR, Hexner EO. A common, under-recognized cause of elevated serum erythropoietin. Am J Med. 2014;127(12):e7-8.
Heterophilic antibodies | Calcitonin (serum) | Falsely elevated calcitonin | Gulbahar O, Konca Degertekin C, Akturk M, Yalcin MM, Kalan I, Atikeler GF, et al. A Case With Immunoassay Interferences in the Measurement of Multiple Hormones. J Clin Endocrinol Metab. 2015;100(6):2147-53.
Heterophilic antibodies | Total PSA (Prostate specific antigen) (serum) | Falsely elevated Total PSA | Fritz BE, Hauke RJ, Stickle DF. New onset of heterophilic antibody interference in prostate-specific antigen measurement occurring during the period of post-prostatectomy prostate-specific antigen monitoring. Ann Clin Biochem. 2009;46(Pt 3):253-6.
Heterophilic antibodies | Ferritin (serum) | Falsely elevated Ferritin | Blanco M, Varela C. Interference from heterophilic antibodies in the Olympus ferritin method. Clin Chem. 2006;52(8):1623-4.
Heterophilic antibodies | Insulin (serum) | Falsely elevated insulin | Craven M, Lord K, Leavens KF, De Leon DD. Insulin Immunoassay Interference Due to Human Antimouse Antibodies in a Patient With Ketotic Hypoglycemia. JCEM Case Rep. 2023;1(2):luad029.
Heterophilic antibodies | 25-OH Vit D (25-hydroxyvitamin D) (serum) | Falsely elevated 25 OH Vitamin D | Hager HB, Bolstad N, Warren DJ, Ness MV, Seierstad B, Lindberg M. Falsely markedly elevated 25-hydroxyvitamin D in patients with monoclonal gammopathies. Clin Chem Lab Med. 2021;59(4):663-9.
Macrohormones | Prolactin (serum) | Falsely elevated Prolactin | Overgaard M, Pedersen SM. Serum prolactin revisited: parametric reference intervals and cross platform evaluation of polyethylene glycol precipitation-based methods for discrimination between hyperprolactinemia and macroprolactinemia. Clin Chem Lab Med. 2017;55(11):1744-53.
Macrohormones | FSH (follicle stimulating hormone) (serum) | Falsely elevated FSH | Webster R, Fahie-Wilson M, Barker P, Chatterjee VK, Halsall DJ. Immunoglobulin interference in serum follicle-stimulating hormone assays: autoimmune and heterophilic antibody interference. Ann Clin Biochem. 2010;47(Pt 4):386-9.
Macrohormones | TSH (Thyroid stimulating hormone) (serum) | Falsely elevated TSH | Chiardi I, Rotondi M, Cantù M, Keller F, Trimboli P. Macro-TSH: An Uncommon Explanation for Persistent TSH Elevation That Thyroidologists Have to Keep in Mind. J Pers Med. 2023;13(10).
Anti-Streptavidin antibodies | FT3 (free triiodothyronine) (serum) | Falsely elevated FT3 | Favresse J, Burlacu MC, Maiter D, Gruson D. Interferences With Thyroid Function Immunoassays: Clinical Implications and Detection Algorithm. Endocr Rev. 2018;39(5):830-50.
Anti-Streptavidin antibodies | FT4 (free thyroxine) (serum) | Falsely elevated FT4 | Favresse J, Burlacu MC, Maiter D, Gruson D. Interferences With Thyroid Function Immunoassays: Clinical Implications and Detection Algorithm. Endocr Rev. 2018;39(5):830-50.
Anti-Streptavidin antibodies | Anti-TG (thyroglobulin antibody) (serum) | Falsely elevated Anti-Tg | Paczkowska K, Otlewska A, Loska O, Kolačkov K, Bolanowski M, Daroszewski J. Laboratory interference in the thyroid function test. Endokrynol Pol. 2020;71(6):551-60.
Anti-Streptavidin antibodies | Anti-TPO (Thyroid peroxidase antibody) (serum) | Falsely elevated Anti-TPO | Paczkowska K, Otlewska A, Loska O, Kolačkov K, Bolanowski M, Daroszewski J. Laboratory interference in the thyroid function test. Endokrynol Pol. 2020;71(6):551-60.
Anti-Streptavidin antibodies | Anti-CCP (cyclic citrullinated peptide antibody) (serum) | Falsely elevated Anti-CCP | Berth M, Willaert S, De Ridder C. Anti-streptavidin IgG antibody interference in anti-cyclic citrullinated peptide (CCP) IgG antibody assays is a rare but important cause of false-positive anti-CCP results. Clin Chem Lab Med. 2018;56(8):1263-8.
Anti-Streptavidin antibodies | TSH (Thyroid stimulating hormone) (serum) | Falsely elevated or decreased TSH | Favresse J, Burlacu MC, Maiter D, Gruson D. Interferences With Thyroid Function Immunoassays: Clinical Implications and Detection Algorithm. Endocr Rev. 2018;39(5):830-50.
Anti-animal antibodies | CRP (C-Reactive Protein) (serum) | Falsely elevated C-reactive protein | Benoist JF, Orbach D, Biou D. False increase in C-reactive protein attributable to heterophilic antibodies in two renal transplant patients treated with rabbit antilymphocyte globulin. Clin Chem. 1998;44(9):1980-5.
Anti-animal antibodies | FSH (follicle stimulating hormone) (serum) | Falsely elevated FSH | Fruzzetti F, Palla G, Sbrana A, Simoncini T, Sessa MR. Anti-goat antibodies as a rare cause of high gonadotropin levels during menopausal transition. Gynecol Endocrinol. 2020;36(10):938-40.
Anti-animal antibodies | LH (Luteinizing hormone) (serum) | Falsely elevated LH | Fruzzetti F, Palla G, Sbrana A, Simoncini T, Sessa MR. Anti-goat antibodies as a rare cause of high gonadotropin levels during menopausal transition. Gynecol Endocrinol. 2020;36(10):938-40.
Anti-animal antibodies | Prolactin (serum) | Falsely elevated Prolactin | Fruzzetti F, Palla G, Sbrana A, Simoncini T, Sessa MR. Anti-goat antibodies as a rare cause of high gonadotropin levels during menopausal transition. Gynecol Endocrinol. 2020;36(10):938-40. | Aliberti L, Gagliardi I, Dorizzi RM, Pizzicotti S, Bondanelli M, Zatelli MC, et al. Hypeprolactinemia: still an insidious diagnosis. Endocrine. 2021;72(3):928-31.
Hook effect | Prolactin (serum) | Falsely decreased Prolactin | Frieze TW, Mong DP, Koops MK. "Hook effect" in prolactinomas: case report and review of literature. Endocr Pract. 2002;8(4):296-303.
Hook effect | B-HCG (Beta-Human chorionic gonadotropin) (serum) | Falsely decreased B-HCG | Cormano J, Mackay G, Holschneider C. Gestational Trophoblastic Disease Diagnosis Delayed by the Hook Effect. Obstet Gynecol. 2015;126(4):811-4.
Anti-ruthenium antibodies | TSH (Thyroid stimulating hormone) (serum) | Falsely elevated, rarely descreaed TSH | Favresse J, Burlacu MC, Maiter D, Gruson D. Interferences With Thyroid Function Immunoassays: Clinical Implications and Detection Algorithm. Endocr Rev. 2018;39(5):830-50.
Anti-ruthenium antibodies | FT3 (free triiodothyronine) (serum) | Falsely elevated, rarely descreaed FT3 | Favresse J, Burlacu MC, Maiter D, Gruson D. Interferences With Thyroid Function Immunoassays: Clinical Implications and Detection Algorithm. Endocr Rev. 2018;39(5):830-50.
Anti-ruthenium antibodies | FT4 (free thyroxine) (serum) | Falsely elevated, rarely descreaed FT4 | Favresse J, Burlacu MC, Maiter D, Gruson D. Interferences With Thyroid Function Immunoassays: Clinical Implications and Detection Algorithm. Endocr Rev. 2018;39(5):830-50.
Anti-Tg antibody | Thyroglobulin (serum) | Falsely decreased Thyroglobulin | Jo K, Lim DJ. Clinical implications of anti-thyroglobulin antibody measurement before surgery in thyroid cancer. Korean J Intern Med. 2018;33(6):1050-7.
Anti-insulin antibody | Insulin (serum) | Falsely decreased insulin | Sapin R. The Interference of Insulin Antibodies in Insulin Immunometric Assays. 2002;40(7):705-8.
Rheumatoid factors (RFs) | Thyroglobulin (serum) | Falsely elevated thyroglobulin | Astarita G, Gutiérrez S, Kogovsek N, Mormandi E, Otero P, Calabrese C, et al. False positive in the measurement of thyroglobulin induced by rheumatoid factor. Clin Chim Acta. 2015;447:43-6.
Rheumatoid factors (RFs) | LH (Luteinizing hormone) (serum) | Falsely elevated LH | Mongolu S, Armston AE, Mozley E, Nasruddin A. Heterophilic antibody interference affecting multiple hormone assays: Is it due to rheumatoid factor? Scand J Clin Lab Invest. 2016;76(3):240-2.
Rheumatoid factors (RFs) | FSH (follicle stimulating hormone) (serum) | Falsely elevated FSH | Mongolu S, Armston AE, Mozley E, Nasruddin A. Heterophilic antibody interference affecting multiple hormone assays: Is it due to rheumatoid factor? Scand J Clin Lab Invest. 2016;76(3):240-2.
Rheumatoid factors (RFs) | SHBG (Sex hormone binding globulin) (serum) | Falsely elevated SHBG | Mongolu S, Armston AE, Mozley E, Nasruddin A. Heterophilic antibody interference affecting multiple hormone assays: Is it due to rheumatoid factor? Scand J Clin Lab Invest. 2016;76(3):240-2.
Rheumatoid factors (RFs) | Prolactin (serum) | Falsely elevated Prolactin | Mongolu S, Armston AE, Mozley E, Nasruddin A. Heterophilic antibody interference affecting multiple hormone assays: Is it due to rheumatoid factor? Scand J Clin Lab Invest. 2016;76(3):240-2.
Rheumatoid factors (RFs) | B-HCG (Beta-Human chorionic gonadotropin) (serum) | Falsely elevated B-HCG | Mongolu S, Armston AE, Mozley E, Nasruddin A. Heterophilic antibody interference affecting multiple hormone assays: Is it due to rheumatoid factor? Scand J Clin Lab Invest. 2016;76(3):240-2.
Rheumatoid factors (RFs) | TSH (Thyroid stimulating hormone) (serum) | Falsely elevated TSH | Mongolu S, Armston AE, Mozley E, Nasruddin A. Heterophilic antibody interference affecting multiple hormone assays: Is it due to rheumatoid factor? Scand J Clin Lab Invest. 2016;76(3):240-2.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | Cortisol (free and total) (serum) | Falsely elevated Cortisol | Hawley JM, Owen LJ, Lockhart SJ, Monaghan PJ, Armston A, Chadwick CA, et al. Serum Cortisol: An Up-To-Date Assessment of Routine Assay Performance. Clin Chem. 2016;62(9):1220-9.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | Estradiol (serum) | Falsely elevated estradiol | Mandic S, Kratzsch J, Mandic D, Debeljak Z, Lukic I, Horvat V, et al. Falsely elevated serum oestradiol due to exemestane therapy. Ann Clin Biochem. 2017;54(3):402-5.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | Growth hormone (serum) | Falsely elevated GH | Ribeiro de Oliveira Longo Schweizer J, Ribeiro-Oliveira Jr A, Bidlingmaier M. Growth hormone: isoforms, clinical aspects and assays interference. Clinical Diabetes and Endocrinology. 2018;4(1):18.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | 11-Deoxycortisol (serum) | Falsely elevated 11-Deoxycortisol | Wu AHB. Tietz Clinical Guide to Laboratory Tests. 4 ed: Saunders/Elsevier; 2006.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | Free Cortisol (in 24-hour urine) | Falsely elevated Free Cortisol (24-hour urine) | Meikle AW, Findling J, Kushnir MM, Rockwood AL, Nelson GJ, Terry AH. Pseudo-Cushing syndrome caused by fenofibrate interference with urinary cortisol assayed by high-performance liquid chromatography. J Clin Endocrinol Metab. 2003;88(8):3521-4.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | Aldosterone (serum) | Falsely elevated aldosterone | Taylor PJ, Cooper DP, Gordon RD, Stowasser M. Measurement of aldosterone in human plasma by semiautomated HPLC-tandem mass spectrometry. Clin Chem. 2009;55(6):1155-62.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | 17-Hydroxyprogesterone (serum) | Falsely elevated 17-Hydroxyprogesterone | Watanabe K, Tsuji-Hosokawa A, Hashimoto A, Konishi K, Ishige N, Yajima H, et al. The High Relevance of 21-Deoxycortisol, (Androstenedione + 17α-Hydroxyprogesterone)/Cortisol, and 11-Deoxycortisol/17α-Hydroxyprogesterone for Newborn Screening of 21-Hydroxylase Deficiency. J Clin Endocrinol Metab. 2022;107(12):3341-52.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | 17-Hydroxypregnanolone (serum) | Falsely elevated 17-Hydroxypregnanolone | Watanabe K, Tsuji-Hosokawa A, Hashimoto A, Konishi K, Ishige N, Yajima H, et al. The High Relevance of 21-Deoxycortisol, (Androstenedione + 17α-Hydroxyprogesterone)/Cortisol, and 11-Deoxycortisol/17α-Hydroxyprogesterone for Newborn Screening of 21-Hydroxylase Deficiency. J Clin Endocrinol Metab. 2022;107(12):3341-52.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | Epinephrine | Falsely elevated Epinephrine | Davidson FD. Paracetamol-associated interference in an HPLC-ECD assay for urinary free metadrenalines and catecholamines. Ann Clin Biochem. 2004;41(Pt 4):316-20.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | Dopamine | Falsely elevated Dopamine | Davidson FD. Paracetamol-associated interference in an HPLC-ECD assay for urinary free metadrenalines and catecholamines. Ann Clin Biochem. 2004;41(Pt 4):316-20.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | Metanephrine | Falsely elevated Metanephrine | Davidson FD. Paracetamol-associated interference in an HPLC-ECD assay for urinary free metadrenalines and catecholamines. Ann Clin Biochem. 2004;41(Pt 4):316-20.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | Normetanephrine | Falsely elevated Normetanephrine | Davidson FD. Paracetamol-associated interference in an HPLC-ECD assay for urinary free metadrenalines and catecholamines. Ann Clin Biochem. 2004;41(Pt 4):316-20.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | Vanilmandelic Acid | Falsely elevated Vanilmandelic Acid | Levreri I, Caruso U, Deiana F, Buoncompagni A, Bernardi BD, Marchese N, et al. The secretion of ibuprofen metabolites interferes with the capillary chromatography of urinary homovanillic acid and 4-hydroxy-3-methoxymandelic acid in neuroblastoma diagnosis. Clinical Chemistry and Laboratory Medicine (CCLM). 2005;43(2):173-7.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | Homovanilic Acid | Falsely elevated Homovanilic Acid | Levreri I, Caruso U, Deiana F, Buoncompagni A, Bernardi BD, Marchese N, et al. The secretion of ibuprofen metabolites interferes with the capillary chromatography of urinary homovanillic acid and 4-hydroxy-3-methoxymandelic acid in neuroblastoma diagnosis. Clinical Chemistry and Laboratory Medicine (CCLM). 2005;43(2):173-7.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | CA 72-4 (cancer antigen 72-4) (serum) | Falsely elevated CA 72-4 | Liang Y, He M, Fan X, Ye W, Yang Z, Zhong R. An abnormal elevation of serum CA72-4 by ganoderma lucidum spore powder. Ann Clin Lab Sci. 2013;43(3):337-40.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | Anti-dsDNA (Double-stranded DNA antibodies) (serum) | Falsely elevated anti-dsDNA | Egner W. The use of laboratory tests in the diagnosis of SLE. J Clin Pathol. 2000;53(6):424-32.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | Total-Testosterone (serum) | Falsely elevated Total Testosterone | Krasowski MD, Drees D, Morris CS, Maakestad J, Blau JL, Ekins S. Cross-reactivity of steroid hormone immunoassays: clinical significance and two-dimensional molecular similarity prediction. BMC Clin Pathol. 2014;14:33.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | Rheumatoid Factor (serum) | Falsely elevated Rheumatoid Factor | Sonntag O, Scholer A. Drug interference in clinical chemistry: recommendation of drugs and their concentrations to be used in drug interference studies. Ann Clin Biochem. 2001;38(Pt 4):376-85.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | Insulin (serum) | Falsely elevated insulin | Dayaldasani A, Rodríguez Espinosa M, Ocón Sánchez P, Pérez Valero V. Cross-reactivity of insulin analogues with three insulin assays. Ann Clin Biochem. 2015;52(Pt 3):312-8.
Cross-reaction (Metabolites or precursors of the analyte, drugs, and other substances) | ACE (angiotensin converting enzyme) (serum) | Falsely decreased ACE (angiotensin converting enzyme) | Szabó AÁ, Enyedi EE, Altorjay IT, Hajnal P, Pintér TB, Mányiné IS, et al. Get reliable laboratory findings – how to recognize the deceptive effects of angiotensin-converting enzyme inhibitor therapy in the laboratory diagnostics of sarcoidosis? Clinical Chemistry and Laboratory Medicine (CCLM). 2024.
Anti-bovine Alkalane phospatase antibodies | Estradiol (serum) | Falsely elevated or decreased estradiol | Maharjan AS, Wyness SP, Ray JA, Willcox TL, Seiter JD, Genzen JR. Detection and characterization of estradiol (E2) and unconjugated estriol (uE3) immunoassay interference due to anti-bovine alkaline phosphatase (ALP) antibodies. Pract Lab Med. 2019;17:e00131.
Anti-bovine Alkalane phospatase antibodies | Unconjugated estriol | Falsely elevated or decreased unconjugated estriol | Maharjan AS, Wyness SP, Ray JA, Willcox TL, Seiter JD, Genzen JR. Detection and characterization of estradiol (E2) and unconjugated estriol (uE3) immunoassay interference due to anti-bovine alkaline phosphatase (ALP) antibodies. Pract Lab Med. 2019;17:e00131.
Monoclonal IgM-Kappa | Homocysteine (total) (serum) | Falsely increased Homocysteine | Charifa A, El-Khoury JM. Hyperhomocysteinemia in a Gammopathy Patient: Malnutrition or Interference? Clin Chem. 2019;65(8):1063-4.
Monoclonal IgM-Kappa | CRP (C-Reactive Protein) (serum) | Falsely increased CRP | Yamada K, Yagihashi A, Ishii S, Tanemura K, Kida T, Watanabe N, et al. Interference with nephelometric assay of C-reactive protein and antistreptolysin-O by monoclonal IgM-kappa from a myeloma patient. Clin Chem. 1997;43(12):2435-7.
Monoclonal IgM-Kappa | ASO (Antisreptolysin-O) (serum) | Falsely increased ASO | Yamada K, Yagihashi A, Ishii S, Tanemura K, Kida T, Watanabe N, et al. Interference with nephelometric assay of C-reactive protein and antistreptolysin-O by monoclonal IgM-kappa from a myeloma patient. Clin Chem. 1997;43(12):2435-7.
IgM antibody response to the E. coli infection | CA 125 (cancer antigen 125) (serum) | Falsely elevated CA 125 | Covinsky M, Laterza O, Pfeifer JD, Farkas-Szallasi T, Scott MG. An IgM lambda antibody to Escherichia coli produces false-positive results in multiple immunometric assays. Clin Chem. 2000;46(8 Pt 1):1157-61.
IgM antibody response to the E. coli infection | B-HCG (Beta-Human chorionic gonadotropin) (serum) | Falsely elevated B-HCG | Covinsky M, Laterza O, Pfeifer JD, Farkas-Szallasi T, Scott MG. An IgM lambda antibody to Escherichia coli produces false-positive results in multiple immunometric assays. Clin Chem. 2000;46(8 Pt 1):1157-61.
IgM antibody response to the E. coli infection | AFP (Alpha feto protein) (serum) | Falsely elevated AFP | Covinsky M, Laterza O, Pfeifer JD, Farkas-Szallasi T, Scott MG. An IgM lambda antibody to Escherichia coli produces false-positive results in multiple immunometric assays. Clin Chem. 2000;46(8 Pt 1):1157-61.
IgM antibody response to the E. coli infection | TSH (Thyroid stimulating hormone) (serum) | Falsely elevated TSH | Covinsky M, Laterza O, Pfeifer JD, Farkas-Szallasi T, Scott MG. An IgM lambda antibody to Escherichia coli produces false-positive results in multiple immunometric assays. Clin Chem. 2000;46(8 Pt 1):1157-61.
Methotrexate, aminopterin, and folinic acid (Leucovorin) intake | Folic acid (serum) | Falsely elevated Folic acid | Young DS, American Association for Clinical C. Effects of drugs on clinical laboratory tests. 5th ed. Washington, DC: AACC Press Washington, DC; 2000.
Prolonged storage before analysis | Anti-diuretic hormone (ADH) (plasma) | Falsely decreased Anti-diuretic hormone (ADH) | Wu AHB. Tietz Clinical Guide to Laboratory Tests. 4 ed: Saunders/Elsevier; 2006.
Late centrifugation after sample collection | Calcitonin (serum) | Falsely decreased Calcitonin | de Carvalho GA, Drobrzenski B. Challenges in Thyroid Function Testing: Interferences and Clinical Implications. Medical Research Archives. 2023;11(8).
Increased blood lead levels | 1-25-Vit D (1-25-dihydroxyvitamin D) (serum) | Falsely decreased 1-25-dihydroxyvitamin D | Wu AHB. Tietz Clinical Guide to Laboratory Tests. 4 ed: Saunders/Elsevier; 2006.
Hemolysis | MCH (Mean corpuscular hemoglobin) (blood) | Falsely elevated MCH | de Jonge G, Dos Santos TL, Cruz BR, Simionatto M, Bittencourt JIM, Krum EA, et al. Interference of in vitro hemolysis complete blood count. J Clin Lab Anal. 2018;32(5):e22396.
Hemolysis | MCHC (Mean corpuscular hemoglobin concentration) (blood) | Falsely elevated MCHC | de Jonge G, Dos Santos TL, Cruz BR, Simionatto M, Bittencourt JIM, Krum EA, et al. Interference of in vitro hemolysis complete blood count. J Clin Lab Anal. 2018;32(5):e22396.
Hemolysis | RDW (erythrocyte distribution width) (blood) | Falsely elevated RDW | de Jonge G, Dos Santos TL, Cruz BR, Simionatto M, Bittencourt JIM, Krum EA, et al. Interference of in vitro hemolysis complete blood count. J Clin Lab Anal. 2018;32(5):e22396.
Hemolysis | Platelet count (blood) | Falsely elevated PLT | de Jonge G, Dos Santos TL, Cruz BR, Simionatto M, Bittencourt JIM, Krum EA, et al. Interference of in vitro hemolysis complete blood count. J Clin Lab Anal. 2018;32(5):e22396.
Hemolysis | MCV (Mean corpuscular volume) (blood) | Falsely decreased MCV | de Jonge G, Dos Santos TL, Cruz BR, Simionatto M, Bittencourt JIM, Krum EA, et al. Interference of in vitro hemolysis complete blood count. J Clin Lab Anal. 2018;32(5):e22396.
Hemolysis | Erythrocyte count (blood) | Falsely decreased Erythrocyte count | de Jonge G, Dos Santos TL, Cruz BR, Simionatto M, Bittencourt JIM, Krum EA, et al. Interference of in vitro hemolysis complete blood count. J Clin Lab Anal. 2018;32(5):e22396.
Hemolysis | Hematocrit (blood) | Falsely decreased Hematocrit | de Jonge G, Dos Santos TL, Cruz BR, Simionatto M, Bittencourt JIM, Krum EA, et al. Interference of in vitro hemolysis complete blood count. J Clin Lab Anal. 2018;32(5):e22396.
Lipemia | Hemoglobin (blood) | Falsely elevated Hemoglobin | Zeng SG, Zeng TT, Jiang H, Wang LL, Tang SQ, Sun YM, et al. A simple, fast correction method of triglyceride interference in blood hemoglobin automated measurement. J Clin Lab Anal. 2013;27(5):341-5.
Lipemia | MCH (Mean corpuscular hemoglobin) (blood) | Falsely elevated MCH | Zeng SG, Zeng TT, Jiang H, Wang LL, Tang SQ, Sun YM, et al. A simple, fast correction method of triglyceride interference in blood hemoglobin automated measurement. J Clin Lab Anal. 2013;27(5):341-5.
Lipemia | MCHC (Mean corpuscular hemoglobin concentration) (blood) | Falsely elevated MCHC | Zeng SG, Zeng TT, Jiang H, Wang LL, Tang SQ, Sun YM, et al. A simple, fast correction method of triglyceride interference in blood hemoglobin automated measurement. J Clin Lab Anal. 2013;27(5):341-5.
Prolonged storage before analysis | MCV (Mean corpuscular volume) (blood) | Falsely increased MCV | De la Salle B. Pre- and postanalytical errors in haematology. Int J Lab Hematol. 2019;41 Suppl 1:170-6. | Gulati G, Uppal G, Gong J. Unreliable Automated Complete Blood Count Results: Causes, Recognition, and Resolution. Ann Lab Med. 2022;42(5):515-30.
Prolonged storage before analysis | MPV (mean platelet volume) (blood) | Falsely increased MPV | De la Salle B. Pre- and postanalytical errors in haematology. Int J Lab Hematol. 2019;41 Suppl 1:170-6.
Large Platelets | MCV (Mean corpuscular volume) (blood) | Falsely decreased MCV | Dasgupta A, Sepulveda JL. Accurate results in the clinical laboratory: A guide to error detection and correction. 2 ed: Elsevier; 2019.
Cold agglutination on RBC | MCV (Mean corpuscular volume) (blood) | Falsely increased MCV | Topic A, Milevoj Kopcinovic L, Bronic A, Pavic M. Effect of cold agglutinins on red blood cell parameters in a trauma patient: a case report. Biochem Med (Zagreb). 2018;28(3):031001.
Leukocyte Aggregation | Leukocyte count (blood) | Falsely decreased Leukocyte | Yang D, Guo X, Chen Y, Xu G. Leukocyte Aggregation In Vitro as a Cause of Pseudoleukopenia. Laboratory Medicine. 2008;39(2):89-91. | Walradth EA. Lymphocyte Aggregation in Low-Grade B-Cell Lymphoma. Lab Med. 2021;52(5):e133-e6.
Leukocyte Aggregation | Lymphocyte (blood) | Falsely decreased Lymphocyte | Walradth EA. Lymphocyte Aggregation in Low-Grade B-Cell Lymphoma. Lab Med. 2021;52(5):e133-e6.
Leukocyte Aggregation | Neutrophil (blood) | Falsely decreased Neutrophil count | Yang D, Guo X, Chen Y, Xu G. Leukocyte Aggregation In Vitro as a Cause of Pseudoleukopenia. Laboratory Medicine. 2008;39(2):89-91.
Severe microcytosis | MPV (mean platelet volume) (blood) | Falsely higher MPV | Gulati G, Uppal G, Gong J. Unreliable Automated Complete Blood Count Results: Causes, Recognition, and Resolution. Ann Lab Med. 2022;42(5):515-30.
High blood glucose ( ≥600 mg/dL, ≥33 mmol/L) | MCV (Mean corpuscular volume) (blood) | Falsely increased MCV | Morse EE, Kalache G, Germino W, Stockwell R. Increased electronic mean corpuscular volume induced by marked hyperglycemia. Ann Clin Lab Sci. 1981;11(2):184-7.
Striking granulation in the neutrophiles and eosinophiles (Due to Maroteaux–Lamy syndrome) | Basophil (blood) | Falsely elevated Basophil | Piva E, Pelloso M, Ciubotaru D, Penello L, Burlina A, Plebani M. The role of automated analyzers in detecting abnormal granulation of leucocytes in lysosomal storage diseases: Maroteaux-Lamy disease. Am J Hematol. 2013;88(6):527.
Large reactive lymphocytes | Monocyte (blood) | Falsely elevated Monocyte | Dasgupta A, Sepulveda JL. Accurate results in the clinical laboratory: A guide to error detection and correction. 2 ed: Elsevier; 2019.
Heinz bodies | Reticulocyte (blood) | Falsely elevated Reticulocyte | Hinchliffe RF. Errors in automated reticulocyte counts due to Heinz bodies. J Clin Pathol. 1993;46(9):878-9.
Prolonged storage (in room temperature that possibly lead to bacterial proliferation) | pH (urine) | Falsely Increased pH | Cook JD, Strauss KA, Caplan YH, Lodico CP, Bush DM. Urine pH: the effects of time and temperature after collection. J Anal Toxicol. 2007;31(8):486-96.
Prolonged storage (in room temperature that possibly lead to bacterial proliferation) | Nitrite (urine) | Falsely Increased nitrite | Riley RS, McPherson RA. Basic Examination of Urine. Henry's Clinical Diagnosis and Management by Laboratory Methods. 23 ed2017. p. 442-80.
Prolonged storage (in room temperature that possibly lead to bacterial proliferation) | Ketone (urine) | Falsely decreased ketone | Simerville JA, Maxted WC, Pahira JJ. Urinalysis: a comprehensive review. Am Fam Physician. 2005;71(6):1153-62.
Prolonged storage (in room temperature that possibly lead to bacterial proliferation) | Creatinine Clearance (serum, urine) | Falsely decreased creatinine clearance | Brunzel NA. Fundamentals of Urine and Body Fluid Analysis. Philadelphia: Elsevier Philadelphia; 2022. Available from: https://public.ebookcentral.proquest.com/choice/PublicFullRecord.aspx?p=7044914.
Alkaline Urine ( from bacterial contamination or alkaline buffer usage) | Protein (urine) | Falsely Increased Protein | Riley RS, McPherson RA. Basic Examination of Urine. Henry's Clinical Diagnosis and Management by Laboratory Methods. 23 ed2017. p. 442-80.
Contamination with oxidizing substance | Glucose (urine) | Falsely Increased glucose (chemical analysis with strip) | Delanghe J, Speeckaert M. Preanalytical requirements of urinalysis. Biochem Med (Zagreb). 2014;24(1):89-104. | Smith BC, Peake MJ, Fraser CG. Urinalysis by use of multi-test reagent strips: two dipsticks compared. Clin Chem. 1977;23(12):2337-40.
Contamination with oxidizing substance | Hemoglobin/erythrocyte (urine with strip analysis) | Falsely Increased blood (chemical analysis with strip) | Delanghe J, Speeckaert M. Preanalytical requirements of urinalysis. Biochem Med (Zagreb). 2014;24(1):89-104. | Smith BC, Peake MJ, Fraser CG. Urinalysis by use of multi-test reagent strips: two dipsticks compared. Clin Chem. 1977;23(12):2337-40.
Microbial Peroxidase (Associated with Urinary Tract Infection) | Hemoglobin/erythrocyte (urine with strip analysis) | Falsely Increased blood (chemical analysis with strip) | Riley RS, McPherson RA. Basic Examination of Urine. Henry's Clinical Diagnosis and Management by Laboratory Methods. 23 ed2017. p. 442-80.
Contamination of urine with vaginal ﬂuid | Leucocyte Esterase (urine) | Falsely Increased Leucocyte Esterase | Riley RS, McPherson RA. Basic Examination of Urine. Henry's Clinical Diagnosis and Management by Laboratory Methods. 23 ed2017. p. 442-80.
Intake of ascorbic acid (generally when ≥25 mg/dL) | Nitrite (urine) | Falsely Decreased Nitrite | Simerville JA, Maxted WC, Pahira JJ. Urinalysis: a comprehensive review. Am Fam Physician. 2005;71(6):1153-62.
Intake of ascorbic acid (generally when ≥25 mg/dL) | Bilirubin (urine) | Falsely Decreased Bilirubin | Brunzel NA. Fundamentals of Urine and Body Fluid Analysis. Philadelphia: Elsevier Philadelphia; 2022. Available from: https://public.ebookcentral.proquest.com/choice/PublicFullRecord.aspx?p=7044914.
Intake of ascorbic acid (generally when ≥25 mg/dL) | Hemoglobin/erythrocyte (urine with strip analysis) | Falsely Decreased erytrocyte | Riley RS, McPherson RA. Basic Examination of Urine. Henry's Clinical Diagnosis and Management by Laboratory Methods. 23 ed2017. p. 442-80.
Light exposure of the sample | Urobilinogen (urine) | Falsely Decreased Urobilinogen | Brunzel NA. Fundamentals of Urine and Body Fluid Analysis. Philadelphia: Elsevier Philadelphia; 2022. Available from: https://public.ebookcentral.proquest.com/choice/PublicFullRecord.aspx?p=7044914.
Light exposure of the sample | Bilirubin (urine) | Falsely Decreased Bilirubin | Riley RS, McPherson RA. Basic Examination of Urine. Henry's Clinical Diagnosis and Management by Laboratory Methods. 23 ed2017. p. 442-80.
Drug and food induced color changes, such as Phenazopyridine (Pyridium), beet | Bilirubin (urine) | Falsely Increased Bilirubin | Simerville JA, Maxted WC, Pahira JJ. Urinalysis: a comprehensive review. Am Fam Physician. 2005;71(6):1153-62.
Drug and food induced color changes, such as Phenazopyridine (Pyridium), beet | Urobilinogen (urine) | Falsely Increased Urobilinogen | Simerville JA, Maxted WC, Pahira JJ. Urinalysis: a comprehensive review. Am Fam Physician. 2005;71(6):1153-62.
Drug and food induced color changes, such as Phenazopyridine (Pyridium), beet | Nitrite (urine) | Falsely Increased nitrite | Riley RS, McPherson RA. Basic Examination of Urine. Henry's Clinical Diagnosis and Management by Laboratory Methods. 23 ed2017. p. 442-80.
Intake of tetracycline, cephalexin, cephalothin | Leucocyte Esterase (urine) | Falsely Decreased Leucocyte Esterase | Strasinger SK, Di Lorenzo MS. Urinalysis and body fluids. Sixth ed. Philadelphia [Pennsylvania], Boston, Massachusetts: F.A. Davis Company, Credo Reference Philadelphia [Pennsylvania], Boston, Massachusetts; 2014.
Intake of Methyldopa, Sodium 2-mercaptoethane sulfonate (MESNA) containing drugs | Ketone (urine) | Falsely Increased Ketone | Riley RS, McPherson RA. Basic Examination of Urine. Henry's Clinical Diagnosis and Management by Laboratory Methods. 23 ed2017. p. 442-80.
Intake of captopril | Ketone (urine) | Falsely Increased Ketone | Riley RS, McPherson RA. Basic Examination of Urine. Henry's Clinical Diagnosis and Management by Laboratory Methods. 23 ed2017. p. 442-80.
Intake of captopril | Hemoglobin/erythrocyte (urine with strip analysis) | Falsely Decreased Eryrtocyte | Riley RS, McPherson RA. Basic Examination of Urine. Henry's Clinical Diagnosis and Management by Laboratory Methods. 23 ed2017. p. 442-80.
Contamination with Formalin | Urobilinogen (urine) | Falsely Decreased Urobilinogen | Brunzel NA. Fundamentals of Urine and Body Fluid Analysis. Philadelphia: Elsevier Philadelphia; 2022. Available from: https://public.ebookcentral.proquest.com/choice/PublicFullRecord.aspx?p=7044914.
Contamination with Formalin | Hemoglobin/erythrocyte (urine with strip analysis) | Falsely Decreased Eryrtocyte | Riley RS, McPherson RA. Basic Examination of Urine. Henry's Clinical Diagnosis and Management by Laboratory Methods. 23 ed2017. p. 442-80.
Nitrite Contamination | Bilirubin (urine) | Falsely Decreased Bilirubin | Riley RS, McPherson RA. Basic Examination of Urine. Henry's Clinical Diagnosis and Management by Laboratory Methods. 23 ed2017. p. 442-80.
Elevated urine speciﬁc gravity, protein, and glucose | Leucocyte Esterase (urine) | Falsely Decreased Leucocyte Esterase | Strasinger SK, Di Lorenzo MS. Urinalysis and body fluids. Sixth ed. Philadelphia [Pennsylvania], Boston, Massachusetts: F.A. Davis Company, Credo Reference Philadelphia [Pennsylvania], Boston, Massachusetts; 2014.
High Ketone Concentration in urine (40 mg/dL) | Glucose (urine) | Falsely Decreased Glucose | Court JM, Davies HE, Ferguson R. Diastix and ketodiastix, a new semiquantitative test for glucose in urine. Med J Aust. 1972;1(11):525-8.
Urine pH<6.0 | Nitrite (urine) | Falsely Decreased Nitrite | Simerville JA, Maxted WC, Pahira JJ. Urinalysis: a comprehensive review. Am Fam Physician. 2005;71(6):1153-62.
Contamination with clorhexidine | Protein (urine) | Falsely Increased Protein | Riley RS, McPherson RA. Basic Examination of Urine. Henry's Clinical Diagnosis and Management by Laboratory Methods. 23 ed2017. p. 442-80.
Contamination of boric acid | Leucocyte Esterase (urine) | Falsely Decreased Leucocyte Esterase | Strasinger SK, Di Lorenzo MS. Urinalysis and body fluids. Sixth ed. Philadelphia [Pennsylvania], Boston, Massachusetts: F.A. Davis Company, Credo Reference Philadelphia [Pennsylvania], Boston, Massachusetts; 2014.
Inclusion of two first morning samples for 24 hour urine collection | Albumin (24 hour urine) | Falsely increased for 24 hour urine albumin | Brunzel NA. Fundamentals of Urine and Body Fluid Analysis. Philadelphia: Elsevier Philadelphia; 2022. Available from: https://public.ebookcentral.proquest.com/choice/PublicFullRecord.aspx?p=7044914.
Inclusion of two first morning samples for 24 hour urine collection | Total Protein (24 hour urine) | Falsely increased for 24 hour urine protein | Brunzel NA. Fundamentals of Urine and Body Fluid Analysis. Philadelphia: Elsevier Philadelphia; 2022. Available from: https://public.ebookcentral.proquest.com/choice/PublicFullRecord.aspx?p=7044914.
Inclusion of two first morning samples for 24 hour urine collection | Citrate (24 hour urine) | Falsely increased for 24 hour urine citrate | Brunzel NA. Fundamentals of Urine and Body Fluid Analysis. Philadelphia: Elsevier Philadelphia; 2022. Available from: https://public.ebookcentral.proquest.com/choice/PublicFullRecord.aspx?p=7044914.
Inclusion of two first morning samples for 24 hour urine collection | Oxalate (24 hour urine) | Falsely increased for 24 hour urine oxalate | Brunzel NA. Fundamentals of Urine and Body Fluid Analysis. Philadelphia: Elsevier Philadelphia; 2022. Available from: https://public.ebookcentral.proquest.com/choice/PublicFullRecord.aspx?p=7044914.
Inaccurate total urine volume measurement | Albumin (24 hour urine) | Falsely decreased or increased 24 hour urine albumin | Brunzel NA. Fundamentals of Urine and Body Fluid Analysis. Philadelphia: Elsevier Philadelphia; 2022. Available from: https://public.ebookcentral.proquest.com/choice/PublicFullRecord.aspx?p=7044914.
Inaccurate total urine volume measurement | Total Protein (24 hour urine) | Falsely decreased or increased 24 hour urine protein | Brunzel NA. Fundamentals of Urine and Body Fluid Analysis. Philadelphia: Elsevier Philadelphia; 2022. Available from: https://public.ebookcentral.proquest.com/choice/PublicFullRecord.aspx?p=7044914.
Inaccurate total urine volume measurement | Citrate (24 hour urine) | Falsely decreased or increased 24 hour urine citrate | Brunzel NA. Fundamentals of Urine and Body Fluid Analysis. Philadelphia: Elsevier Philadelphia; 2022. Available from: https://public.ebookcentral.proquest.com/choice/PublicFullRecord.aspx?p=7044914.
Inaccurate total urine volume measurement | Oxalate (24 hour urine) | Falsely decreased or increased 24 hour urine oxalate | Brunzel NA. Fundamentals of Urine and Body Fluid Analysis. Philadelphia: Elsevier Philadelphia; 2022. Available from: https://public.ebookcentral.proquest.com/choice/PublicFullRecord.aspx?p=7044914.
Delayed examination | Leucocytes (urine) | False negative urine white blood cells | Kierkegaard H, Feldt-Rasmussen U, Hørder M, Andersen HJ, Jørgensen PJ. Falsely negative urinary leucocyte counts due to delayed examination. Scand J Clin Lab Invest. 1980;40(3):259-61.
Delayed examination | Erythrocytes (urine) | False negative urine red blood cells | Dolscheid-Pommerich RC, Klarmann-Schulz U, Conrad R, Stoffel-Wagner B, Zur B. Evaluation of the appropriate time period between sampling and analyzing for automated urinalysis. Biochem Med (Zagreb). 2016;26(1):82-9.
Cells and cell fragments | Yeast (urine) | False negative or positive yeast | Aydin O, Ellidag HY, Eren E, Yilmaz N. High False Positives and False Negatives in Yeast Parameter in an Automated Urine Sediment Analyzer. J Med Biochem. 2015;34(3):332-7.
Urine preservatives | Erythrocytes (urine) | False negative red blood cells | Delanghe J, Speeckaert M. Preanalytical requirements of urinalysis. Biochem Med (Zagreb). 2014;24(1):89-104.
Urine preservatives | Epithelial cells (urine) | False negative epithelial cells | Delanghe J, Speeckaert M. Preanalytical requirements of urinalysis. Biochem Med (Zagreb). 2014;24(1):89-104.
Urine preservatives | Casts (urine) | False negative casts | Delanghe J, Speeckaert M. Preanalytical requirements of urinalysis. Biochem Med (Zagreb). 2014;24(1):89-104.
Diluted/Hypoosmolar urine | Leucocytes (urine) | False negative urine white blood cells | Patel HP. The abnormal urinalysis. Pediatr Clin North Am. 2006;53(3):325-37, v.
Diluted/Hypoosmolar urine | Erythrocytes (urine) | False negative urine red blood cells | Patel HP. The abnormal urinalysis. Pediatr Clin North Am. 2006;53(3):325-37, v.
Alkaline pH | Leucocytes (urine) | False negative urine white blood cells | Fogazzi GB, Verdesca S, Garigali G. Urinalysis: core curriculum 2008. Am J Kidney Dis. 2008;51(6):1052-67.
Alkaline pH | Erythrocytes (urine) | False negative urine red blood cells | Fogazzi GB, Verdesca S, Garigali G. Urinalysis: core curriculum 2008. Am J Kidney Dis. 2008;51(6):1052-67.
Insufficient sample volume | PT (prothrombin time) (plasma) | Falsely prolongs PT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Insufficient sample volume | aPTT (activated partial thromboplastin time) (plasma) | Falsely prolongs aPTT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Insufficient sample volume | D-dimer (plasma) | Falsely decreased D-dimer | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Insufficient sample volume | Fibrinogen (plasma) | Falsely decreased fibrinogen | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Hyperlipidemia | PT (prothrombin time) (plasma) | Falsely shortened PT (Optic metodes) | Dasgupta A, Sepulveda JL. Accurate results in the clinical laboratory: A guide to error detection and correction. 2 ed: Elsevier; 2019.
Hyperlipidemia | aPTT (activated partial thromboplastin time) (plasma) | Falsely shortened aPTT (Optic metodes) | Dasgupta A, Sepulveda JL. Accurate results in the clinical laboratory: A guide to error detection and correction. 2 ed: Elsevier; 2019.
Hyperbilirubinemia | PT (prothrombin time) (plasma) | Falsely shortened PT (Optic metodes) | Dasgupta A, Sepulveda JL. Accurate results in the clinical laboratory: A guide to error detection and correction. 2 ed: Elsevier; 2019.
Hyperbilirubinemia | aPTT (activated partial thromboplastin time) (plasma) | Falsely shortened aPTT (Optic metodes) | Dasgupta A, Sepulveda JL. Accurate results in the clinical laboratory: A guide to error detection and correction. 2 ed: Elsevier; 2019.
Serum or fully clotted sample | PT (prothrombin time) (plasma) | Falsely shortened PT | Favaloro EJ, Funk DM, Lippi G. Pre-analytical Variables in Coagulation Testing Associated With Diagnostic Errors in Hemostasis. Laboratory Medicine. 2012;43(2):1-10.
Serum or fully clotted sample | aPTT (activated partial thromboplastin time) (plasma) | Falsely shortened aPTT | Favaloro EJ, Funk DM, Lippi G. Pre-analytical Variables in Coagulation Testing Associated With Diagnostic Errors in Hemostasis. Laboratory Medicine. 2012;43(2):1-10.
Serum or fully clotted sample | Thrombin Time (TT) (plasma) | Falsely shortened TT | Favaloro EJ, Funk DM, Lippi G. Pre-analytical Variables in Coagulation Testing Associated With Diagnostic Errors in Hemostasis. Laboratory Medicine. 2012;43(2):1-10.
Serum or fully clotted sample | D-dimer (plasma) | Falsely decreased D-Dimer | Favaloro EJ, Funk DM, Lippi G. Pre-analytical Variables in Coagulation Testing Associated With Diagnostic Errors in Hemostasis. Laboratory Medicine. 2012;43(2):1-10.
Inappropriate sample (Usage of EDTA as anticoagulant) | PT (prothrombin time) (plasma) | Falsely increased PT | Favaloro EJ, Funk DM, Lippi G. Pre-analytical Variables in Coagulation Testing Associated With Diagnostic Errors in Hemostasis. Laboratory Medicine. 2012;43(2):1-10.
Inappropriate sample (Usage of EDTA as anticoagulant) | aPTT (activated partial thromboplastin time) (plasma) | Falsely increased aPTT | Favaloro EJ, Funk DM, Lippi G. Pre-analytical Variables in Coagulation Testing Associated With Diagnostic Errors in Hemostasis. Laboratory Medicine. 2012;43(2):1-10.
Inappropriate sample (Usage of filtered plasma for transfusion purpose) | Fibrinogen (plasma) | Loss of fibrinogen | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Inappropriate sample (Usage of filtered plasma for transfusion purpose) | PT (prothrombin time) (plasma) | Falsely prolonged PT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Inappropriate sample (Usage of filtered plasma for transfusion purpose) | aPTT (activated partial thromboplastin time) (plasma) | Falsely prolonged aPTT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Inappropriate sample (Usage of filtered plasma for transfusion purpose) | Thrombin Time (TT) (plasma) | Falsely prolonged TT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Inappropriate sample (Usage of sodium-heparin as anticoagulant) | aPTT (activated partial thromboplastin time) (plasma) | No clotting in aPTT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Inappropriate sample (Usage of sodium-heparin as anticoagulant) | PT (prothrombin time) (plasma) | No clotting in PT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Inappropriate sample (Usage of sodium-heparin as anticoagulant) | Thrombin Time (TT) (plasma) | No clotting in TT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Inappropriate sample (Usage of sodium-heparin as anticoagulant) | Protein C activity (plasma) | Falsely decreased Protein C activity | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Inappropriate sample (Usage of sodium-heparin as anticoagulant) | Protein S activity (plasma) | Falsely decreased Protein S activity | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Inappropriate sample (Usage of sodium-heparin as anticoagulant) | anti-Xa | Falsely increased anti-Xa | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Heparin contamination (ex-vivo or due to collection tube error) | PT (prothrombin time) (plasma) | Falsely prolonged PT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Heparin contamination (ex-vivo or due to collection tube error) | aPTT (activated partial thromboplastin time) (plasma) | Falsely prolonged aPTT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Heparin contamination (ex-vivo or due to collection tube error) | Thrombin Time (TT) (plasma) | Falsely prolonged TT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Heparin contamination (ex-vivo or due to collection tube error) | Fibrinogen (plasma) | Falsely decreased fibrinogen | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Heparin contamination (ex-vivo or due to collection tube error) | Antithrombin (plasma) | Falsely decreased antithrombin | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Delayed transport or testing | PT (prothrombin time) (plasma) | Falsely prolonged PT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Delayed transport or testing | aPTT (activated partial thromboplastin time) (plasma) | Falsely prolonged aPTT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Delayed transport or testing | aPTT (Unfractionated Heparinized Samples) | Falsely Shortened aPTT (Unfractionated Heparinized Samples) | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Delayed transport or testing | anti-FXa (heparin) test | Falsely decreased anti-FXa (heparin) test | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Several Freeze- Thaw Events | PT (prothrombin time) (plasma) | Falsely prolonged PT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Several Freeze- Thaw Events | aPTT (activated partial thromboplastin time) (plasma) | Falsely prolonged aPTT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Several Freeze- Thaw Events | aPTT (Unfractionated Heparinized Samples) | Falsely shortened aPTT (for unfractionated heparinized Samples) | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Several Freeze- Thaw Events | anti-FXa (heparin) test | Falsely decreased anti-FXa (heparin) test | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Storage in Frost-Free Freezer | PT (prothrombin time) (plasma) | Falsely prolonged PT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Storage in Frost-Free Freezer | aPTT (activated partial thromboplastin time) (plasma) | Falsely prolonged aPTT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Poor centrifugation | aPTT (activated partial thromboplastin time) (plasma) | Falsely shortened aPTT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Poor centrifugation | Anti-FXa (heparin) test | Falsely decreased anti-FXa (heparin) test | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Poor centrifugation | Lupus anticoagulant (plasma) | Falsely decreased LA | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Heavy braking of centrifuge | aPTT (activated partial thromboplastin time) (plasma) | Falsely shortened aPTT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Heavy braking of centrifuge | Anti-FXa (heparin) test | Falsely decreased anti-FXa (heparin) test | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Heavy braking of centrifuge | Lupus anticoagulant (plasma) | Falsely decreased LA | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Sample remixing prior to freezing | aPTT (activated partial thromboplastin time) (plasma) | Falsely shortened aPTT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Sample remixing prior to freezing | Anti-FXa (heparin) test | Falsely decreased anti-FXa (heparin) test | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Sample remixing prior to freezing | Lupus anticoagulant (plasma) | Falsely decreased LA | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Antithrombin inhibitor anticoagulant intake | PT (prothrombin time) (plasma) | Falsely prolonged PT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Antithrombin inhibitor anticoagulant intake | aPTT (activated partial thromboplastin time) (plasma) | Falsely prolonged aPTT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Antithrombin inhibitor anticoagulant intake | Thrombin Time (TT) (plasma) | Falsely prolonged TT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Antithrombin inhibitor anticoagulant intake | Fibrinogen (plasma) | Falsely decreased fibrinogen | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Antithrombin inhibitor anticoagulant intake | Protein C (clot-based) (plasma) | Falsely increased Protein C (for clot-based assay) | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Antithrombin inhibitor anticoagulant intake | Protein S (clot-based) (plasma) | Falsely increased Protein S (for clot-based assay) | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Anti-Xa inhibitor anticoagulant intake | PT (prothrombin time) (plasma) | Falsely prolonged PT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Anti-Xa inhibitor anticoagulant intake | aPTT (activated partial thromboplastin time) (plasma) | Falsely prolonged aPTT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Anti-Xa inhibitor anticoagulant intake | Thrombin Time (TT) (plasma) | Falsely prolonged TT | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Anti-Xa inhibitor anticoagulant intake | Antithrombin (Xa-based assays) (plasma) | Falsely increased antithrombin (Xa-based assays) | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Anti-Xa inhibitor anticoagulant intake | Protein C (clot-based) (plasma) | Falsely increased Protein C (for clot-based assay) | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Anti-Xa inhibitor anticoagulant intake | Protein S (clot-based) (plasma) | Falsely increased Protein S (for clot-based assay) | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Anti-Xa inhibitor anticoagulant intake | Lupus anticoagulant (plasma) | False negative LA | Adcock DM, Favaloro EJ, Lippi G. Critical pre-examination variables in the hemostasis laboratory and their quality indicators. Clin Biochem. 2016;49(18):1315-20.
Heterophilic antibodies | D-dimer (plasma) | Falsely increased D Dimer | Verboogen D, Granzen B, Hudig C, van de Kerkhof D, Verhezen P, de Boer D, et al. Heterophilic antibodies leading to falsely positive D-dimer concentration in an adolescent. Res Pract Thromb Haemost. 2023;7(1):100017.
Increased Acute Phase Reactans | aPTT (activated partial thromboplastin time) (plasma) | Falsely shortened aPTT | Dasgupta A, Sepulveda JL. Accurate results in the clinical laboratory: A guide to error detection and correction. 2 ed: Elsevier; 2019.
Rheumatoid factors (RFs) | D-dimer (plasma) | Falsely increased D-Dimer | Dasgupta A, Sepulveda JL. Accurate results in the clinical laboratory: A guide to error detection and correction. 2 ed: Elsevier; 2019.
Digoxin Like Immunreactive Factors | Digoxin (serum) | Falsely increased digoxin | Sánchez BM, De Cos MA, Peralta FG, Arribas C, Armijo JA. Syva Emit 2000 and Roche "on line" subject to less interference by digoxin-like factors than Abbott TDx FPIA in newborns and pregnant women. Clinical Chemistry. 1996;42(6):974-6.
Fab products (DIGIBAND and Digifab) | Digoxin (serum) | Falsely increased digoxin | Dasgupta A, Wahed A. Chapter 14 - Therapeutic drug monitoring. In: Dasgupta A, Wahed A, editors. Clinical Chemistry, Immunology and Laboratory Quality Control (Second Edition): Elsevier; 2021. p. 295-323.
Fab products (DIGIBAND and Digifab) | Digoxin (serum) | Falsely decreased digoxin | Dasgupta A, Wahed A. Chapter 14 - Therapeutic drug monitoring. In: Dasgupta A, Wahed A, editors. Clinical Chemistry, Immunology and Laboratory Quality Control (Second Edition): Elsevier; 2021. p. 295-323.
Cardiac glycoside intake | Digoxin (serum) | Falsely increased digoxin | Dasgupta A, Scott j. Unexpected Suppression of Total Digoxin Concentrations by Cross-Reactants in the Microparticle Enzyme Immunoassay: Elimination of Interference by Monitoring Free Digoxin Concentration. American Journal of Clinical Pathology. 1998;110(1):78-82.
Cardiac glycoside intake | Digoxin (serum) | Falsely decreased digoxin | Dasgupta A, Scott j. Unexpected Suppression of Total Digoxin Concentrations by Cross-Reactants in the Microparticle Enzyme Immunoassay: Elimination of Interference by Monitoring Free Digoxin Concentration. American Journal of Clinical Pathology. 1998;110(1):78-82.
Aldosteron antagonist intake | Digoxin (serum) | Falsely increased or decreased digoxin | Steimer W, Müller C, Eber B. Digoxin assays: frequent, substantial, and potentially dangerous interference by spironolactone, canrenone, and other steroids. Clin Chem. 2002;48(3):507-16.
Lithium Heparin Tubes | Lithium (serum) | Falsely increased lithium | Dasgupta A, Wahed A. Chapter 14 - Therapeutic drug monitoring. In: Dasgupta A, Wahed A, editors. Clinical Chemistry, Immunology and Laboratory Quality Control (Second Edition): Elsevier; 2021. p. 295-323.
Nonsteroidal anti-inflammatory drugs and diuretics (Thiazid) intake | Lithium (serum) | Falsely increased lithium | Nunes RP. Lithium interactions with non-steroidal anti-inflammatory drugs and diuretics – A review. Archives of Clinical Psychiatry (São Paulo). 2018;45.
Hyperbilirubinemia | Tacrolimus (blood) | Falsely increased tacrolimus | Gonschior AK, Christians U, Winkler M, Linck A, Baumann J, Sewing KF. Tacrolimus (FK506) metabolite patterns in blood from liver and kidney transplant patients. Clin Chem. 1996;42(9):1426-32.
Endogenous anti-ß-galactosidase antibodies | Tacrolimus (blood) | Falsely increased tacrolimus | D'Alessandro M, Mariani P, Mennini G, Severi D, Berloco P, Bachetoni A. Falsely elevated tacrolimus concentrations measured using the ACMIA method due to circulating endogenous antibodies in a kidney transplant recipient. Clin Chim Acta. 2011;412(3-4):245-8.
Endogenous anti-ß-galactosidase antibodies | Cyclosporine (blood) | Falsely increased cyclosporin | D'Alessandro M, Mariani P, Mennini G, Severi D, Berloco P, Bachetoni A. Falsely elevated tacrolimus concentrations measured using the ACMIA method due to circulating endogenous antibodies in a kidney transplant recipient. Clin Chim Acta. 2011;412(3-4):245-8.
Cross-reaction with 4-deoxy-4-amino-N10-methylpteroic acid (DAMPA) | Methotrexate (serum) | Falsely increased methotrexate levels | Oudart JB, Marquet B, Feliu C, Gozalo C, Djerada Z, Millart H. Analytical interference in the therapeutic drug monitoring of methotrexate. Ann Biol Clin (Paris). 2016;74(3):333-7.
Prolonged storage in room temperature | Valproic acid (serum) | Falsely increased valproic acid | Tarasidis CG, Garnett WR, Kline BJ, Pellock JM. Influence of tube type, storage time, and temperature on the total and free concentration of valproic acid. Ther Drug Monit. 1986;8(3):373-6.
Serum seperator tube gel absorption | Valproic acid (serum) | Falsely decreased valproic acid | Warner A, Privitera M, Bates D. Standards of laboratory practice: antiepileptic drug monitoring. National Academy of Clinical Biochemistry. Clin Chem. 1998;44(5):1085-95.
Serum seperator tube gel absorption | Phenytoin (serum) | Falsely decreased phenytoine | Clinical and Laboratory Standards Institute. Procedures for the Handling and Processing of Blood Specimens for Common Laboratory Tests; Approved Guideline—Fourth Edition. CLSI document GP44-A4. Wayne, Pennsylvania 19087 USA: Clinical and Laboratory Standards Institute; 2010.
Cross-reaction with Ibuprofen and Naproxen | Barbiturate (urine) | False positive urine barbiturate | Saitman A, Park HD, Fitzgerald RL. False-positive interferences of common urine drug screen immunoassays: a review. J Anal Toxicol. 2014;38(7):387-96. | Rollins DE, Jennison TA, Jones G. Investigation of interference by nonsteroidal anti-inflammatory drugs in urine tests for abused drugs. Clin Chem. 1990;36(4):602-6.
Oxaprozin and 5-(p-hydroxyphenyl)-5-phenylhydantoin interference | Phenytoin (serum) | Falsely increased serum phenytoine | Datta P. Oxaprozin and 5-(p-hydroxyphenyl)-5-phenylhydantoin interference in phenytoin immunoassays. Clin Chem. 1997;43(8 Pt 1):1468-9. | Datta P, Dasgupta A. Cross-reactivity of fosphenytoin in four phenytoin immunoassays. Clin Chem. 1998;44(3):696-7.
IgM lambda interference | Phenytoin (serum) | Falsely decreased serum phenytoine | Brauchli YB, Scholer A, Schwietert M, Krähenbühl S. Undetectable phenytoin serum levels by an automated particle-enhanced turbidimetric inhibition immunoassay in a patient with monoclonal IgM lambda. Clin Chim Acta. 2008;389(1-2):174-6.
Cross-reactivity with Efavirenz and Sertraline | Benzodiazepine (urine) | Falsely positive result in urine benzodiazepine | Saitman A, Park HD, Fitzgerald RL. False-positive interferences of common urine drug screen immunoassays: a review. J Anal Toxicol. 2014;38(7):387-96.
Cross-reactions with 10,11-epoxide metabolites. | Carbamazepine (serum) | Falsely increased carbamazepine | Hermida J, Tutor JC. How suitable are currently used carbamazepine immunoassays for quantifying carbamazepine-10,11-epoxide in serum samples? Ther Drug Monit. 2003;25(3):384-8.
St. John’s wort interference | Theophylline (serum) | Falsely decreased theophylline | Izzo AA, Ernst E. Interactions between herbal medicines and prescribed drugs: an updated systematic review. Drugs. 2009;69(13):1777-98.
Wrong specimen (plasma instead serum ) | Protein electrophoresis (serum) | Misinterpreted as a monoclonal immunoglobulin (among β and γ region) | Plebani M, Banfi G, Bernardini S, Bondanini F, Conti L, Dorizzi R, et al. Serum or plasma? An old question looking for new answers. Clin Chem Lab Med. 2020;58(2):178-87.
Piperacillin-tazobactam interference | Protein electrophoresis (serum) | A small peak at the anodal site of the β-globulin fraction | Bossuyt X, Peetermans WE. Effect of Piperacillin-Tazobactam on Clinical Capillary Zone Electrophoresis of Serum Proteins. Clinical Chemistry. 2002;48(1):204-5.
Elevated transferrin levels interference | Protein electrophoresis (serum) | A strong band in the beta region | Dasgupta A, Wahed A. Chapter 24 - Protein electrophoresis and immunofixation. In: Dasgupta A, Wahed A, editors. Clinical Chemistry, Immunology and Laboratory Quality Control (Second Edition): Elsevier; 2021. p. 489-511.
Intravascular hemolysis | Protein electrophoresis (serum) | A band in the alpha-2 region | Dasgupta A, Wahed A. Chapter 24 - Protein electrophoresis and immunofixation. In: Dasgupta A, Wahed A, editors. Clinical Chemistry, Immunology and Laboratory Quality Control (Second Edition): Elsevier; 2021. p. 489-511.
Aggregation of paraproteins | Protein electrophoresis (serum) | A broad smear band | Dasgupta A, Wahed A. Chapter 24 - Protein electrophoresis and immunofixation. In: Dasgupta A, Wahed A, editors. Clinical Chemistry, Immunology and Laboratory Quality Control (Second Edition): Elsevier; 2021. p. 489-511.
Rapid urine elimination in cases of light chain myeloma | Immunoglobulin free light chains (kappa and lambda) (serum) | Falsely negative light chain result | Dasgupta A, Wahed A. Chapter 24 - Protein electrophoresis and immunofixation. In: Dasgupta A, Wahed A, editors. Clinical Chemistry, Immunology and Laboratory Quality Control (Second Edition): Elsevier; 2021. p. 489-511.
Hemolysis | Trace elements panel | Falsely increased trace element | Laur N, Kinscherf R, Pomytkin K, Kaiser L, Knes O, Deigner HP. ICP-MS trace element analysis in serum and whole blood. PLoS One. 2020;15(5):e0233357.
Sample contamination | Trace elements panel | Falsely increased trace element | Versieck J, Barbier F, Cornelis R, Hoste J. Sample contamination as a source of error in trace-element analysis of biological samples. Talanta. 1982;29(11 Pt 2):973-84.
Contamination during sample collection and preparation | Trace elements panel | Falsely increased trace element | Versieck J, Barbier F, Cornelis R, Hoste J. Sample contamination as a source of error in trace-element analysis of biological samples. Talanta. 1982;29(11 Pt 2):973-84. | Versieck J, Vanballenberghe L, De Kesel A, Van Renterghem D. Accuracy of biological trace-element determinations. Biol Trace Elem Res. 1987;12(1):45-54.
Gadolinium contrast interference | Selenium (serum) | Falsely increased selenium | Ryan JB, Grant S, Walmsley T, Florkowski CM, George PM. Falsely elevated plasma selenium due to gadolinium contrast interference: a novel solution to a preanalytical problem. Ann Clin Biochem. 2014;51(Pt 6):714-6.
Total parenteral nutrition solutions interference | Amino acids (plasma) | Falsely increased amino asids levels | Chace DH, De Jesús VR, Lim TH, Hannon WH, Clark RH, Spitzer AR. Detection of TPN contamination of dried blood spots used in newborn and metabolic screening and its impact on quantitative measurement of amino acids. Clin Chim Acta. 2011;412(15-16):1385-90.
Ethylene glycol intoxication | Lactate (plasma) | Falsely increased lactate levels | Sandberg Y, Rood PP, Russcher H, Zwaans JJ, Weige JD, van Daele PL. Falsely elevated lactate in severe ethylene glycol intoxication. Neth J Med. 2010;68(1):320-3.
Inappropriate phlebotomy | Lactate (plasma) | Falsely increased lactate levels | Patel K, Jones PM. Specimen collection and processing. In: Rifai N, editor. Tietz Textbook of Laboratory Medicine. 7th ed: Elsevier; 2023. p. 61-79.
Prolonged contact of plasma and serum with cells at room temperature | Lactate (plasma) | Falsely increased lactate levels | Boyanton BL, Jr., Blick KE. Stability studies of twenty-four analytes in human plasma and serum. Clin Chem. 2002;48(12):2242-7.
Prolonged sample storage | Lactate (plasma) | Falsely increased lactate levels | Astles R, Williams CP, Sedor F. Stability of plasma lactate in vitro in the presence of antiglycolytic agents. Clin Chem. 1994;40(7 Pt 1):1327-30.
Delayed sample processing | Ammonia (plasma) | Falsely increased ammonia levels | Maranda B, Cousineau J, Allard P, Lambert M. False positives in plasma ammonia measurement and their clinical impact in a pediatric population. Clin Biochem. 2007;40(8):531-5.
Cigarette smoke contamination | Ammonia (plasma) | Falsely increased ammonia levels | Willems EW, Rambali B, Vleeming W, Opperhuizen A, van Amsterdam JGC. Significance of ammonium compounds on nicotine exposure to cigarette smokers. Food and Chemical Toxicology. 2006;44(5):678-88.
Prolonged sample storage before analysis | Ammonia (plasma) | Falsely increased ammonia levels | Hashim IA, Cuthbert JA. Elevated ammonia concentrations: potential for pre-analytical and analytical contributing factors. Clin Biochem. 2014;47(16-17):233-6.
Prolonged tourniquet time | pH (blood) | Falsely decreased pH | Cengiz M, Ulker P, Meiselman HJ, Baskurt OK. Influence of tourniquet application on venous blood sampling for serum chemistry, hematological parameters, leukocyte activation and erythrocyte mechanical properties. Clin Chem Lab Med. 2009;47(6):769-76.
Delayed analysis and storage on ice in plastic syringe | pO2 (partial pressure of oxygen) (blood) | Falsely decreased or increased pO2 | Beaulieu M, Lapointe Y, Vinet B. Stability of PO2, PCO2, and pH in fresh blood samples stored in a plastic syringe with low heparin in relation to various blood-gas and hematological parameters. Clin Biochem. 1999;32(2):101-7.
Interference caused by anaesthetic gases (halothane, nitrous oxide, isoflurane) | pO2 (partial pressure of oxygen) (blood) | Falsely increased pO2 | Eberhard P, Mindt W. Interference of anesthetic gases at oxygen sensors. Birth Defects Orig Artic Ser. 1979;15(4):65-74.
Delayed analysis and storage at room temperature in gas-impermeable containers | pH (blood) | Falsely decreased pH | Beaulieu M, Lapointe Y, Vinet B. Stability of PO2, PCO2, and pH in fresh blood samples stored in a plastic syringe with low heparin in relation to various blood-gas and hematological parameters. Clin Biochem. 1999;32(2):101-7.
Delayed analysis and storage at room temperature in gas-impermeable containers | pO2 (partial pressure of oxygen) (blood) | Falsely decreased pO2 | Beaulieu M, Lapointe Y, Vinet B. Stability of PO2, PCO2, and pH in fresh blood samples stored in a plastic syringe with low heparin in relation to various blood-gas and hematological parameters. Clin Biochem. 1999;32(2):101-7.
Delayed analysis and storage at room temperature in gas-impermeable containers | pCO2 (partial pressure of carbon dioxide) (blood) | Falsely increased pCO2 | Beaulieu M, Lapointe Y, Vinet B. Stability of PO2, PCO2, and pH in fresh blood samples stored in a plastic syringe with low heparin in relation to various blood-gas and hematological parameters. Clin Biochem. 1999;32(2):101-7.
Sampling through a butterfly infusion set | pO2 (partial pressure of oxygen) (blood) | Falsely increased pO2 | Thelin OP, Karanth S, Pourcyrous M, Cooke RJ. Overestimation of neonatal PO2 by collection of arterial blood gas values with the butterfly infusion set. J Perinatol. 1993;13(1):65-7.
Inadequate blood collection into a syringe washed with liquid heparin or washed with too much liquid heparin | pCO2 (partial pressure of carbon dioxide) (blood) | Falsely decreased pCO2 | Küme T, Sişman AR, Solak A, Tuğlu B, Cinkooğlu B, Coker C. The effects of different syringe volume, needle size and sample volume on blood gas analysis in syringes washed with heparin. Biochem Med (Zagreb). 2012;22(2):189-201. | Baird G. Preanalytical considerations in blood gas analysis. Biochem Med (Zagreb). 2013;23(1):19-27.
Inadequate blood collection into a syringe washed with liquid heparin or washed with too much liquid heparin | Bicarbonate (HCO3-) (blood) | Falsely decreased Bicarbonate | Baird G. Preanalytical considerations in blood gas analysis. Biochem Med (Zagreb). 2013;23(1):19-27.
Inadequate blood collection into a syringe washed with liquid heparin or washed with too much liquid heparin | Ionized calcium (blood) | Falsely decreased Ionized calcium | Küme T, Sişman AR, Solak A, Tuğlu B, Cinkooğlu B, Coker C. The effects of different syringe volume, needle size and sample volume on blood gas analysis in syringes washed with heparin. Biochem Med (Zagreb). 2012;22(2):189-201.
Inadequate blood collection into a syringe washed with liquid heparin or washed with too much liquid heparin | pO2 (partial pressure of oxygen) (blood) | Falsely increased pO2 | Küme T, Sişman AR, Solak A, Tuğlu B, Cinkooğlu B, Coker C. The effects of different syringe volume, needle size and sample volume on blood gas analysis in syringes washed with heparin. Biochem Med (Zagreb). 2012;22(2):189-201. | Baird G. Preanalytical considerations in blood gas analysis. Biochem Med (Zagreb). 2013;23(1):19-27.
Transporting samples through pneumatic tube | pO2 (partial pressure of oxygen) (blood) | Falsely decreased pO2 | Collinson PO, John CM, Gaze DC, Ferrigan LF, Cramp DG. Changes in blood gas samples produced by a pneumatic tube system. J Clin Pathol. 2002;55(2):105-7.
Leukocytosis/ thrombocytosis | pO2 (partial pressure of oxygen) (blood) | Falsely decreased pO2 | Schmidt C, Müller-Plathe O. Stability of pO2, pCO2 and pH in heparinized whole blood samples: influence of storage temperature with regard to leukocyte count and syringe material. Eur J Clin Chem Clin Biochem. 1992;30(11):767-73.
Air aspiration or bubble formation | pCO2 (partial pressure of carbon dioxide) (blood) | Falsely decreased pCO2 | Lu JY, Kao JT, Chien TI, Lee TF, Tsai KS. Effects of air bubbles and tube transportation on blood oxygen tension in arterial blood gas analysis. J Formos Med Assoc. 2003;102(4):246-9.
Air aspiration or bubble formation | pH (blood) | Falsely increased pH | Lu JY, Kao JT, Chien TI, Lee TF, Tsai KS. Effects of air bubbles and tube transportation on blood oxygen tension in arterial blood gas analysis. J Formos Med Assoc. 2003;102(4):246-9.
Air aspiration or bubble formation | pO2 (partial pressure of oxygen) (blood) | Falsely increased pO2 | Lu JY, Kao JT, Chien TI, Lee TF, Tsai KS. Effects of air bubbles and tube transportation on blood oxygen tension in arterial blood gas analysis. J Formos Med Assoc. 2003;102(4):246-9.
Air aspiration or bubble formation | Oxygen saturation (sO2) (blood) | Falsely increased Oxygen saturation | Lu JY, Kao JT, Chien TI, Lee TF, Tsai KS. Effects of air bubbles and tube transportation on blood oxygen tension in arterial blood gas analysis. J Formos Med Assoc. 2003;102(4):246-9.
Aberrant hemoglobin or fetal hemoglobin | Oxygen saturation (sO2) (blood) | Falsely decreased or increased Oxygen saturation | Verhovsek M, Henderson MP, Cox G, Luo HY, Steinberg MH, Chui DH. Unexpectedly low pulse oximetry measurements associated with variant hemoglobins: a systematic review. Am J Hematol. 2010;85(11):882-5.
Miscalculation caused by conditions that change the oxygen dissociation curve (presence of fetal hemoglobin, dyshemoglobins or hemoglobinopathies, changes in 2,3-diphosphoglycerate concentrations) | Oxygen saturation (sO2) (blood) | Falsely decreased or increased oxygen saturation (for calculated sO2) | Haymond S, Cariappa R, Eby CS, Scott MG. Laboratory assessment of oxygenation in methemoglobinemia. Clin Chem. 2005;51(2):434-44.
Spectral interference caused by hydroxycobalamine or methylene blue | Oxygen saturation (sO2) (blood) | Falsely decreased or increased oxygen saturation | Clinical and Laboratory Standards Institute. Blood Gas and pH Analysis and Related Measurements; Approved Guideline—Second Edition. CLSI document C46-A2. Wayne, PA: Clinical and Laboratory Standards Institute; 2009.
Delayed umbilical cord clamping | pH (cord blood) | Falsely decreased pH (in cord blood) | Wiberg N, Källén K, Olofsson P. Delayed umbilical cord clamping at birth has effects on arterial and venous blood gases and lactate concentrations. Bjog. 2008;115(6):697-703.
Delayed umbilical cord clamping | Bicarbonate (HCO3-) (blood) | Falsely decreased bicarbonate (in cord blood) | Wiberg N, Källén K, Olofsson P. Delayed umbilical cord clamping at birth has effects on arterial and venous blood gases and lactate concentrations. Bjog. 2008;115(6):697-703.
Delayed umbilical cord clamping | Base deficit (cBase(Ecf)) (blood) | Falsely decreased Base deficit (in cord blood) | Wiberg N, Källén K, Olofsson P. Delayed umbilical cord clamping at birth has effects on arterial and venous blood gases and lactate concentrations. Bjog. 2008;115(6):697-703.
Delayed umbilical cord clamping | pCO2 (partial pressure of carbon dioxide) (blood) | Falsely increased pCO2 (in cord blood) | Wiberg N, Källén K, Olofsson P. Delayed umbilical cord clamping at birth has effects on arterial and venous blood gases and lactate concentrations. Bjog. 2008;115(6):697-703.
Delayed umbilical cord clamping | pO2 (partial pressure of oxygen) (blood) | Falsely increased pO2 (in cord blood) | Wiberg N, Källén K, Olofsson P. Delayed umbilical cord clamping at birth has effects on arterial and venous blood gases and lactate concentrations. Bjog. 2008;115(6):697-703.
Delayed umbilical cord clamping | Lactate (blood) | Falsely increased lactate (in cord blood) | Wiberg N, Källén K, Olofsson P. Delayed umbilical cord clamping at birth has effects on arterial and venous blood gases and lactate concentrations. Bjog. 2008;115(6):697-703.
Spectral interference caused by HbM variants | Methemoglobin (blood) | Falsely decreased methemoglobin | Haymond S, Cariappa R, Eby CS, Scott MG. Laboratory assessment of oxygenation in methemoglobinemia. Clin Chem. 2005;51(2):434-44.
Spectral interference caused by HbM variants | Carboxyhemoglobin (COHb) (blood) | Falsely increased carboxyhemoglobin | Haymond S, Cariappa R, Eby CS, Scott MG. Laboratory assessment of oxygenation in methemoglobinemia. Clin Chem. 2005;51(2):434-44.
Spectral interference caused by HbM variants | Sulfhemoglobin (blood) | Falsely increased sulfhemoglobin | Haymond S, Cariappa R, Eby CS, Scott MG. Laboratory assessment of oxygenation in methemoglobinemia. Clin Chem. 2005;51(2):434-44.
Spectral interference caused by high doses of hydroxycobalamine | Carboxyhemoglobin (COHb) (blood) | Falsely decreased carboxyhemoglobin | Fueyo L, Robles J, Aguilar I, Yáñez AM, Socias M, Parera M. Hemolysis index to detect degree of hydroxocobalamin interference with common laboratory tests. J Clin Lab Anal. 2017;31(5).
Cross reaction, lack of specificity of lactate oxidase and reaction with ethylene glycol metabolites | Lactate (blood) | Falsely increased lactate | Brindley PG, Butler MS, Cembrowski G, Brindley DN. Falsely elevated point-of-care lactate measurement after ingestion of ethylene glycol. Cmaj. 2007;176(8):1097-9.
Underfilling the syringe | Lactate (blood) | Falsely increased lactate | Lippi G, Pighi L, Tosi M, Vettori M, Celegon G, Favaloro EJ, et al. Effect of syringe underfilling on the quality of venous blood gas analysis. Diagnosis (Berl). 2024;11(1):91-6.
Spectral interference caused by sulfhemoglobin | Methemoglobin (blood) | Falsely increased methemoglobin | Zoppi F, Brenna S, Fumagalli C, Marocchi A. Discrimination among dyshemoglobins: analytical approach to a toxicological query. Clinical Chemistry. 1996;42(8):1300-2.
Spectral interference caused by methylene blue | Oxyhemoglobin (blood) | Falsely increased oxyhemoglobin | Gourlain H, Buneaux F, Borron SW, Gouget B, Levillain P. Interference of methylene blue with CO-Oximetry of hemoglobin derivatives. Clin Chem. 1997;43(6 Pt 1):1078-80.
Relative alkalinity of deoxygenated hemoglobin (Haldane effect) | Base deficit (cBase(Ecf)) (blood) | Falsely decreased or increased base deficit | Mentel A, Bach F, Schüler J, Herrmann W, Koster A, Crystal GJ, et al. Assessing errors in the determination of base excess. Anesth Analg. 2002;94(5):1141-8, table of contents.
Variability of the constant coefficients in the Hendersone-Hesselbach equation | Bicarbonate (HCO3-) (blood) | Falsely decreased or increased bicarbonate | Natelson S, Nobel D. Effect of the variation of pK' of the Henderson-Hasselbalch equation on values obtained for total CO2 calculated from pCO2 and pH values. Clin Chem. 1977;23(4):767-9.
Incorrect pregnancy duration calculation, incorrect or incomplete filling of request forms, incorrectly evaluated nuchal translucency | Prenatal screening panel | False positive prenatal screening result | Anumba DO. Errors in prenatal diagnosis. Best Pract Res Clin Obstet Gynaecol. 2013;27(4):537-48.
Hemolysis | Prenatal screening panel | False positive prenatal screening result | Guo Y, Yu D, Zhou K, Wang J, Lei D, Xu Z, et al. The effect of hemolysis on quality control metrics for noninvasive prenatal testing. BMC Med Genomics. 2022;15(1):125.
Prolonged sample storage | Prenatal screening panel | False positive prenatal screening result | Lambert-Messerlian GM, Eklund EE, Malone FD, Palomaki GE, Canick JA, D'Alton ME. Stability of first- and second-trimester serum markers after storage and shipment. Prenat Diagn. 2006;26(1):17-21.
Hemolysis | Retinol binding protein (RBP) (serum) | Falsely decreased Retinol binding protein | Phipps WS, Greene DN, Pflaum H, Laha TJ, Dickerson JA, Irvine J, et al. Small volume retinol binding protein measurement by liquid chromatography-tandem mass spectrometry. Clin Biochem. 2022;99:111-7.
Hemolysis | Tissue transglutaminase antibodies, Ig A and IgG (serum) | Falsely increased Tissue transglutaminase antibodies | Van Meensel B, Hiele M, Hoffman I, Vermeire S, Rutgeerts P, Geboes K, et al. Diagnostic accuracy of ten second-generation (human) tissue transglutaminase antibody assays in celiac disease. Clin Chem. 2004;50(11):2125-35.
Lipemia | Complement C3 | Falsely elevated C3 | Agarwal S, Vargas G, Nordstrom C, Tam E, Buffone GJ, Devaraj S. Effect of interference from hemolysis, icterus and lipemia on routine pediatric clinical chemistry assays. Clin Chim Acta. 2015;438:241-5.
Lipemia | Complement C4 | Falsely elevated C4 | Agarwal S, Vargas G, Nordstrom C, Tam E, Buffone GJ, Devaraj S. Effect of interference from hemolysis, icterus and lipemia on routine pediatric clinical chemistry assays. Clin Chim Acta. 2015;438:241-5.
Lipemia | Haptoglobin (serum) | Falsely elevated Haptoglobulin | Agarwal S, Vargas G, Nordstrom C, Tam E, Buffone GJ, Devaraj S. Effect of interference from hemolysis, icterus and lipemia on routine pediatric clinical chemistry assays. Clin Chim Acta. 2015;438:241-5.
Lipemia | Retinol binding protein (RBP) (serum) | Falsely elevated Retinol binding protein | Phipps WS, Greene DN, Pflaum H, Laha TJ, Dickerson JA, Irvine J, et al. Small volume retinol binding protein measurement by liquid chromatography-tandem mass spectrometry. Clin Biochem. 2022;99:111-7.
Lipemia | Immunoglobulin A (serum) | Falsely elevated Immunoglobulin A | Agarwal S, Vargas G, Nordstrom C, Tam E, Buffone GJ, Devaraj S. Effect of interference from hemolysis, icterus and lipemia on routine pediatric clinical chemistry assays. Clin Chim Acta. 2015;438:241-5. | Bossuyt X, Blanckaert N. Evaluation of interferences in rate and fixed-time nephelometric assays of specific serum proteins. Clin Chem. 1999;45(1):62-7.
Lipemia | Immunoglobulin G (serum) | Falsely elevated Immunoglobulin G | Agarwal S, Vargas G, Nordstrom C, Tam E, Buffone GJ, Devaraj S. Effect of interference from hemolysis, icterus and lipemia on routine pediatric clinical chemistry assays. Clin Chim Acta. 2015;438:241-5.
Lipemia | Alpha-1 antitrypsin (serum) | Falsely elevated Alpha-1 antitrypsin | Ji JZ, Meng QH. Evaluation of the interference of hemoglobin, bilirubin, and lipids on Roche Cobas 6000 assays. Clin Chim Acta. 2011;412(17-18):1550-3.
Lipemia | Tissue transglutaminase antibodies (serum) | Falsely elevated Tissue transglutaminase antibodies | Van Meensel B, Hiele M, Hoffman I, Vermeire S, Rutgeerts P, Geboes K, et al. Diagnostic accuracy of ten second-generation (human) tissue transglutaminase antibody assays in celiac disease. Clin Chem. 2004;50(11):2125-35.
Lipemia | Immunoglobulin M (serum) | Falsely increased or decreased Immunoglobulin M | Bossuyt X, Blanckaert N. Evaluation of interferences in rate and fixed-time nephelometric assays of specific serum proteins. Clin Chem. 1999;45(1):62-7. | Steen G, Vermeer HJ, Naus AJ, Goevaerts B, Agricola PT, Schoenmakers CH. Multicenter evaluation of the interference of hemoglobin, bilirubin and lipids on Synchron LX-20 assays. Clin Chem Lab Med. 2006;44(4):413-9.
Lipemia | Ceruloplasmin (serum) | Falsely increased or decreased ceruloplasmin | Ji JZ, Meng QH. Evaluation of the interference of hemoglobin, bilirubin, and lipids on Roche Cobas 6000 assays. Clin Chim Acta. 2011;412(17-18):1550-3. | Bornhorst JA, Roberts RF, Roberts WL. Assay-specific differences in lipemic interference in native and intralipid-supplemented samples. Clin Chem. 2004;50(11):2197-201.
Excessive turbidity | Alpha-2 macroglobulin (serum) | Falsely increased or decreased Alpha-2 macroglobulin | Wu AHB. Tietz Clinical Guide to Laboratory Tests. 4 ed: Saunders/Elsevier; 2006.
Excessive turbidity | Beta-2 microglobulin (serum) | Falsely increased or decreased Beta-2 microglobulin | Wu AHB. Tietz Clinical Guide to Laboratory Tests. 4 ed: Saunders/Elsevier; 2006.
Antirabbit antibodies interference | Haptoglobin (serum) | Falsely increased Haptoglobin | Butch AW. Falsely increased values for rabbit immunoglobulin-based nephelometric immunoassays attributable to human anti-rabbit antibodies. Clin Chem. 2000;46(8 Pt 1):1207-8.
Immunoglobulin polimerization | Immunoglobulin free light chains (serum) | Falsely increased Immunoglobulin free light chains | de Kat Angelino CM, Raymakers R, Teunesen MA, Jacobs JF, Klasen IS. Overestimation of serum kappa free light chain concentration by immunonephelometry. Clin Chem. 2010;56(7):1188-90.
Immunoglobulin cross reaction | Tissue transglutaminase antibodies (serum) | Falsely increased Tissue transglutaminase antibodies | Villalta D, Crovatto M, Stella S, Tonutti E, Tozzoli R, Bizzaro N. False positive reactions for IgA and IgG anti-tissue transglutaminase antibodies in liver cirrhosis are common and method-dependent. Clin Chim Acta. 2005;356(1-2):102-9.
Cross reaction caused by RF | Cardiolipin antibodies (serum) | Falsely increased Cardiolipin antibodies | Lakos G, Teodorescu M. IgM, but not IgA rheumatoid factor interferes with anti-cardiolipin and antiβ2 glycoprotein I measurements: a quantitative analysis. Lupus. 2011;20(6):614-9.
Enzymatic interference with Fusidic acid | Bile acids | Falsely increased Bile acids | Hansson P, Hedström SA, Hultberg B, Ohlin M. Fusidic acid interferes in enzymatic determination of bile acids. Clin Chim Acta. 1982;125(2):241-3.